##### QADATA Summary – July 2004

The Inorganic Blind Sample Project (IBSP) reviews 108 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 79 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-control samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 12/01/03 through 06/24/04; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples that are known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report to help provide a clearer perspective of the overall concern.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY04charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Arsenic, Filtered, ICP-MS, Labcode 2503, Parmcode 1000D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	4

Uranium, Filtered, ICP-MS, Labcode 1797, Parmcode 22703G. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Calcium, Filtered, ICP, Labcode 0659, Parmcode 0915D  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Sodium, Filtered, ICP, Labcode 0675, Parmcode 0930C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

OWQRL

Analyte

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	8

Thallium, Whole-Water Recoverable, GFAA, Labcode 3131, Parmcode 1059(0)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	9

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Filtered ICP-MS labcode 2503 parmcode 1000D | A dramatic shift in the data is evident starting mid-May 2004. | Arsenic bias was traceable to the correction factor algorithm and the likely source was in the software. This error caused arsenic values to be biased high. The solution used to check the accuracy of the correction equation did not show the error. | The software was reinstalled by the instrument vendor, thus restoring the correct correction factor algorithm. A hard disk diagnostic check on the instrument computers was performed and no problems were found with disk memory. All data affected by this problem is in the process of being updated and a RapiNote is being drafted. The correction check solution has been changed to include a known amount of arsenic to verify the correction factor is correcting the target arsenic concentration. The SOP will be amended, and training will be documented. We anticipate having the documents and training completed by July 30, 2004. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | The long-running slight positive bias (6%) is still evident. This bias is not demonstrated by any other of the five analytical methods analyzing for cadmium. | Internal QC and NIST, both third party checks, indicate a bias of about 2%, which is within the acceptance criteria for calibration standards. Concentration of internal QC and NIST checks has been verified by alternate methods. An internal QC review of the other five methods also shows a 2% or less bias, again within the SOP criteria. Interferences are not a cause for the high bias. | The stock standards are due for the annual replacement at this time. New standards are on order. The process for preparation and review of new standards will follow. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary May 2004 | As of March 2004, results indicate the positive bias is less than 5%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL continues to closely monitor analysis. |
| From QADATA Summary March 2004 | Data from January 2004 through March exhibit a slight positive bias (6%). Almost all results are greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | Due to the on-going monitoring of internal QC, a positive bias is still not evident on NWQL QC charts. To verify this, an additional third-party check standard from a different source than is routinely used by ICP-MS was run during the week of March 8th, 2004, to verify the ICP-MS calibration curve. Calibration and QC are acceptable. | The issue of possible interferences that could affect cadmium is being discussed with the NWQL’s Method Research and Development Program. More research needs to be done. Also, internal QC and IBSP charts will continue to be monitored for bias and variability. |
| From QADATA Summary January 2004 | Data for the last six months (June 2003 through December) demonstrate a slight positive bias (6%) with all 48 values greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of December 2003. We will continue to monitor internal QC and IBSP charts for bias and variability. If the third-party check standards indicate a bias, calibration standards will be remade and verified. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | The negative bias demonstrated in March and April 2004 is no longer evident; however, the most recent data show a slight positive bias (5%). The overall bias since December 2003 is less than 2%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Previous comments for chromium, whole-water recoverable, GFAA, labcode 1555, parmcode 1027F |
| From QADATA Summary May 2004 | Data received as of March 2004 indicate a negative bias (8.5%). | The bias is not evident on NWQL QC charts. Calibration standards have been checked for accuracy and have been verified. Study is being conducted to establish if any instrument modifications should be performed. | NWQL continues to closely monitor analysis for bias and evaluate the need for instrument modifications. |

| Uranium, Filtered ICP-MS labcode 1797 parmcode 22703G | Results received as of mid-April 2004 show no overall bias, and all values are within +/-1 f-pseudosigma. No corrective action necessary. |  |  |
| --- | --- | --- | --- |
| Previous comments for uranium, filtered, ICP-MS, labcode 1797, parmcode 22703G |
| From QADATA Summary May 2004 | As of March 2004 variability has increased. Four out of 15 values (27%) are at or outside of control limits (+/-2 f-pseudosigmas) as of March 2004. | NWQL QC charts show increased variability. | New uranium standards were purchased, prepared, and verified by April 19th. Since then, NWQL QC Charts show reduced variability. The NWQL will continue to closely monitor the analysis for variability. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered ICP labcode 0659 parmcode 0915D Magnesium, Filtered ICP labcode 0663 parmcode 0925C Sodium, Filtered ICP labcode 0675 parmcode 0930C | Data received since the last QADATA Summary indicate an improvement in the bias for each analyte (all less than 4%). The variability has improved for calcium and is now within acceptable levels; however, magnesium and sodium data still exhibit high variability. | Based on internal QC and instrument optimization procedures the variability mentioned by IBSP is not observed. Contamination, calibration curve design, nebulizer conditions and performance, and sample stability for the longer analytical times have all been examined and proven not to be contributors to the variability concerns. The instrument meets the manufacturers specifications for performance. The method meets or exceeds the requirements for ASTM, Standard Methods, and EPA published methods for performance. | We have nothing new to try at this time. However, we will continue to consider the 4% variability and try to make improvements. |
| Previous comments for calcium, magnesium, and sodium by ICP |
| From QADATA Summary May 2004 | Data received since the last QADATA Summary indicate a positive bias. In addition, more values have been received that are at or outside of +/-2 f-pseudosigmas. Results from the new instruments (put into service June 2003) still indicate high variability relative to results from the old instruments. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last QADATA Summary (March 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. Tests have been performed that rule out non-instrument related contamination. Tests are being conducted to determine whether instrument contamination and carry-over is present. The positive bias for calcium and magnesium on BQS control charts is also evident on NWQL QC charts, prompting the preparation of new standards. The positive bias for sodium is not evident on NWQL control charts. | NWQL continues to investigate the cause of the outliers on runs with acceptable QC by running tests on the instrument and determining whether contamination and carry-over is present. New standards for calcium, magnesium, and sodium have been prepared and were validated on May 12th, 2004. NWQL will continue to closely monitor the analysis for bias. |
| From QADATA Summary March 2004 | New instrumentation was implemented in June 2003. A comparison of results from the old instruments (TJA) to results from the new instruments (PE) indicates an increase in variability. The OWQRL is not demonstrating variability problems for these constituents. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last data summary (January 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. | As stated in the January 2004 QADATA Summary for cadmium WWR GFAA (labcode 1555), use of an additional QC check sheet that is included in all data packets was implemented to help alleviate incorrect data going in LIMS. This practice was implemented for ICP processes in October 2003, and data shows that it has been useful in eliminating some errors. We will continue to investigate the cause of the outliers on runs with acceptable QC. |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | The most recently received data show marked improvement. Due to the pending closure of the OWQRL, no response is required. |  |  |
| Previous comments for arsenic, whole-water recoverable, GFAA, labcode 3119, parmcode 1002(1) |
| From QADATA Summary May 2004 | The most recently received data indicate that this method is still demonstrating variability beyond acceptable levels. Of the last seven values received, one has been within the control limits. | Due to the pending closure of the laboratory, new standards were not ordered. | Current stock was used for new standards. |
| From QADATA Summary March 2004 | Since October 2002, data for this analytical line have demonstrated variability beyond acceptable levels. Since October 2002, nine out of 28 points (32%) are outside of +/-2 f-pseudosigmas with three of 28 (11%) outside of +/-3 f-pseudosigmas. The most recently received data indicate continuing problems with the method variability. IBSP recommends that the source of the problem be determined before releasing any more data from environmental samples. | The most recent problems seem to come from the standards used for arsenic. The standard mixes used for ICP have been used for GFAA analyses instead of the standards normally used for GFAA. The arsenic standard was also found not to be matrix matched for the digestion. | New standards for the GFAA will be ordered, checked and matrix matched for whole-water recoverable samples. Any samples requesting arsenic that we have will then be rerun. |
| From QADATA Summary January 2004 | The data still indicate high variability. Two of the last seven values received (29%) are outside of +/-2 f-pseudosigmas. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| From QADATA Summary Nov 2003 | High variability is still demonstrated. Three of the last seven values received (43%) are outside of +/-2 f-pseudosigmas. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| From QADATA Summary Sept 2003 | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 f-pseudosigmas with 3 of 19 (16%) outside of +/-3 f-pseudosigmas. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| Thallium, Whole-Water Recoverable GFAA labcode 3131 parmcode 1059(0) | Data received as of mid-March show no bias, and all values are within +/-1 f-pseudosigma. Due to the pending closure of the OWQRL, no response is required. |  |  |
| Previous comments for thallium, whole-water recoverable, GFAA, labcode 3131, parmcode 1059(0) |
| From QADATA Summary May 2004 | As of January 2004, results indicate a negative bias (22%). This bias is not evident for other methods analyzing for thallium. | Problems with thallium standards. Due to pending closure of the laboratory, new standards were not ordered. | Current stock was used for new standards. |