##### QADATA Summary – March 2004

The Inorganic Blind Sample Project (IBSP) reviews 108 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 79 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-control samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 08/01/03 through 03/04/04; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-control charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have deteriorated or are known to have been affected by a problem at IBSP are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report to help provide a clearer perspective of the overall concern.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY03_04charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Arsenic, Filtered, GFAA, Labcode 2160, Parmcode 1000C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Beryllium, Whole-Water Recoverable, ICP, Labcode 2353, Parmcode 1012C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Filtered , ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Chromium, Filtered, GFAA, Labcode 1936 Parmcode 1030I . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	5

Calcium, Filtered, ICP, Labcode 0659, Parmcode 0915D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Sodium, Filtered, ICP, Labcode 0675, Parmcode 0930C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Residue on Evaporation, Filtered, Gravimetric, Labcode 0027, Parmcode 70300A . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

OWQRL

Analyte

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	7

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Filtered GFAA labcode 2160 parmcode 1000C | This analytical line is still demonstrating a slight positive bias (6.1%). Nineteen out of 20 values are greater than their median value. | The slight positive bias is also evident on QC charts for third-party check samples (TPC). New standards were made and maintenance was performed during the week of March 8th, 2004. | After making new standards and performing maintenance, the positive bias is no longer evident. This line will continue to be monitored for bias. |
| Previous comments for arsenic, filtered, GFAA, labcode 2160, parmcode 1000C |
| From QADATA Summary January 2004 | Data since October 2003 indicate a slight positive bias (7%). Since October there are 11 values (all results) greater than their median value. | New standards were made in November of 2003 including Continuing Calibration Verifications (CCV) and Third-Party Check standards (TPC). Both the CCV’s and TPC’s for arsenic are showing a positive bias on QC charts. | New calibration standards will be remade and verified as of February 6th, 2004. The element will continue to be monitored closely for a bias. |
| Beryllium, Whole-Water Recoverable ICP labcode 2353 parmcode 1012C | Due to the low number of requests for this test, it is no longer monitored by IBSP. |  |  |
| Previous comments for beryllium, whole-water recoverable, ICP, labcode 2353, parmcode 1012C |
| From QADATA Summary January 2004 | New ICP instrumentation came online around July 2003. Prior to July 2003, the data show no bias. Since July 2003, data demonstrate a slight negative bias (8%). | New standards were made in November 2003. Since then, QC charts for CCV’s and TPC’s show a slight positive bias. The six BQS samples were rerun and showed a slight positive bias. Before the standards were made in November, Beryllium was showing a slightly high bias and the standards were nearing their expiration date, thus, necessitating the need for new calibration standards. | New calibration standard will be remade and verified by February 2nd, 2004. The element will continue to be monitored closely for a bias. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | Data from January 2004 through March exhibit a slight positive bias (6%). Almost all results are greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | Due to the on-going monitoring of internal QC, a positive bias is still not evident on NWQL QC charts. To verify this, an additional TPC from a different source than is routinely used by ICP-MS was run during the week of March 8th, 2004, to verify the ICP-MS calibration curve. Calibration and QC are acceptable. | The issue of possible interferences that could affect cadmium is being discussed with the NWQL’s Method Research and Development Program. More research needs to be done. Also, internal QC and IBSP charts will continue to be monitored for bias and variability. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary January 2004 | Data for the last six months (June 2003 through December) demonstrate a slight positive bias (6%) with all 48 values greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of December 2003. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | Data received since the last QADATA Summary suggest the variability has improved. All new data (6 values) are within +/-1 f-pseudosigma of their median value. No corrective action necessary. |  |  |
| Previous comments for cadmium, whole-water recoverable, GFAA, labcode 1555, parmcode 1027F |
| From QADATA Summary January 2004 | As of July 2003 the positive bias has improved (<2%); however, the data now exhibit high variability. Since September there are three out of 11 points (27%) outside of +/-6 f-pseudosigma. | For one of the samples, a blank was accidentally sent to the LIMS instead of a sample and a zero went in the database. There is no obvious explanation as to why the other two points are outside +/- 6 f-pseudosigma. The QC surrounding these samples was in control and the instrument was running within specifications. | To help alleviate human error when putting data into the LIMS, the GFAA section has implemented the use of an additional QC check sheet that is included in all data packets. This sheet lists all samples that have to be rerun due to QC failures, dilutions, possible mispours, and bottles that cannot be found. The analyst initials each sample when the action is performed in LIMS. This should help alleviate the problem of incorrect data going into LIMS. In addition, the section will closely monitor internal QC and IBSP charts for bias and variability. |
| From QADATA Summary November 2003 | A positive bias is still evident; however, the bias has decreased from 7.8 to 5.8%. This same pattern is shown for cadmium, filtered, by GFAA (labcode 1554). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of July. NIST-traceable third party check standards (TPC) were implemented in September. There is not enough data collected from the TPC standards for evaluation. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| From QADATA Summary September 2003 | The results for February 2003 through May 2003 demonstrate a positive bias (7.8%). The results for the same reference materials analyzed at the OWQRL do not exhibit a bias. | Began using new standards July 9, 2003. Online QC charts prior to July 9 show some bias, but bias has decreased since new standards were made. | Improvement is noted since standards were made. NWQL will continue to monitor data quality. |
| From QADATA Summary July 2003 | As of January 2003, a high bias has developed and variability increased. | NWQL QC charts show variable but acceptable data. | Instrument has recently been optimized, and the NWQL will continue to monitor. |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030I | Since late-January 2004, variability has increased and a negative bias has (13.5%) developed. | This line had three samples that fell outside (low) of the +/-2 f-pseudosigma which contributed to the large negative bias. One, with a target value of 6.12, was reported at 1.45. The other two values were between 2 and 3 f-pseudosigma. The negative bias on BQS control charts is not evident on NWQL QC charts. In fact, a slightly high bias was evident on NWQL QC charts at the end of February. | Maintenance was performed on the instrument during the week of March 8, 2004, and early evidence suggests that the positive bias has decreased. Also, the three samples mentioned in the comments section were rerun and were within acceptable limits. The recalculated bias is 4.5%. Analyst will continue to monitor bias and variability and perform maintenance as needed to alleviate both of these problems. NOTE: Three environmental samples associated with the set that had the low value of 1.45 were updated in LIMS and released to the customer. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered ICP labcode 0659 parmcode 00915D Magnesium, Filtered ICP labcode 0663 parmcode 00925C Sodium, Filtered ICP labcode 0675 parmcode 00930C | New instrumentation was implemented in June 2003. A comparison of results from the old instruments (TJA) to results from the new instruments (PE) indicates an increase in variability. The OWQRL is not demonstrating variability problems for these constituents. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last data summary (January 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. | As stated in the January 2004 QADATA Summary for cadmium WWR GFAA (labcode 1555), use of an additional QC check sheet that is included in all data packets was implemented to help alleviate incorrect data going in LIMS. This practice was implemented for ICP processes in October 2003, and data shows that it has been useful in eliminating some errors. We will continue to investigate the cause of the outliers on runs with acceptable QC. |
| Residue on Evaporation, Filtered Gravimetric labcode 0027 parmcode 70300A | A slight negative bias (5%) has developed starting December 2003. | Between December 2003 and February 2004 the NWQL also noticed a low bias with our on-line QC for M-162. The blank and M-152 on-line QC also show a very slight low bias; however, the bias is not as evident. The bias went away near the beginning of March 2004, according to the M-162 on-line QC. All on-line QC are around the expected means as of March 2004. During this time the analyst reported problems with the balance weight being sent to the computer in a consistent manner. The balance doors were adjusted and the timing routine was modified. | The balance communication will continue to be monitored using on-line QC for a reoccurrence of the low bias |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | Since October 2002, data for this analytical line have demonstrated variability beyond acceptable levels. Since October 2002, nine out of 28 points (32%) are outside of +/-2 f-pseudosigma with three of 28 (11%) outside of +/-3 f-pseudosigma. The most recently received data indicate continuing problems with the method variability. IBSP recommends that the source of the problem be determined before releasing any more data from environmental samples. | The most recent problems seem to come from the standards used for arsenic. The standard mixes used for ICP have been used for GFAA analyses instead of the standards normally used for GFAA. The arsenic standard was also found not to be matrix matched for the digestion. | New standards for the GFAA will be ordered, checked and matrix matched for whole-water recoverable samples. Any samples requesting arsenic that we have will then be rerun. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode 1002(1) |
| From QADATA Summary January 2004 | The data still indicate high variability. Two of the last seven values received (29%) are outside of +/-2 f-pseudosigma. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| From QADATA Summary Nov 2003 | High variability is still demonstrated. Three of the last seven values received (43%) are outside of +/-2 f-pseudosigma. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| From QADATA Summary Sept 2003 | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 f-pseudosigma with 3 of 19 (16%) outside of +/-3 f- pseudosigma. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 f-pseudosigma with 3 of 13 (23%) outside of +/-3 f-pseudosigma. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 f-pseudosigma. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |