##### QADATA Summary – February 2005

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems (BQS) reviews about 110 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 07/01/04 through 01/27/05; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ environmental data in a similar manner.  Data from samples known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments from previous QADATA Summaries.  About one month after this report is distributed, the IBSP and the NWQL will meet to follow-up on corrective actions and to discuss new data generated since this report.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY04_05charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Orthophosphorus, Filtered, Colorimetric, Labcode 1974, Parmcode 0671H	4

Chromium, Filtered, GFAA, Labcode 1936, Parmcode 1030I	4

Chromium, Filtered, ICP, Labcode 0722, Parmcode 1030E	5

Copper, Filtered, GFAA, Labcode 1558, Parmcode 1040F	5

Copper, Whole-Water Recoverable, GFAA, Labcode 1559, Parmcode 1042F	5

Copper, Whole-Water Recoverable, ICP-MS, Labcode 2379, Parmcode 1042H	5

Manganese, Filtered, ICP-MS, Labcode 1793, Parmcode 1056G	6

Manganese, Whole-Water Recoverable, ICP-MS, Labcode 2382, Parmcode 1055C	6

Strontium, Filtered, ICP-MS, Labcode 2507, Parmcode 1080I	7

Chloride, Filtered, IC, Labcode 1571, Parmcode 0940J	8

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C	8

Potassium, Filtered, ICP, Labcode 2773, Parmcode 0935C	8

Silica, Filtered, ICP, Labcode 0667, Parmcode 0955D	9

Sulfate, Filtered, IC, Labcode 1572, Parmcode 0945G	9

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Chromium by ICP, filtered, lc 0722 – Positive bias starting November 2004.

- Copper by GFAA, filtered and whole-water recoverable, lcs 1558 and 1559 – Positive bias starting around May 2004.

- Manganese by ICP-MS, filtered and whole-water recoverable, lcs 1793 and 2382 – Negative bias (about 8%) for about six months.

- Potassium by ICP, filtered, lc 2773 – Slight positive bias (5.7%) starting November 2004.

- Strontium by ICP-MS, filtered, lc 2507 – Very high variability.

- Sulfate by IC, filtered, lc 1571 – Negative bias (7.2%) starting November 2004.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Nutrient Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Orthophosphorus, Filtered colorimetric labcode 1974 parmcode 0671H | A negative bias is demonstrated for the latter half of November 2004. Results since then (except one value) exhibit little bias and acceptable variability. No corrective action necessary. | An instrument problem that has been corrected was the reason for the three low data points in late November 2004. |  |

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030I | The positive bias starting July 2004 and going through September is no longer evident. Results as of October 2004 show no measurable bias. No corrective action necessary. |  |  |
| Previous comments for chromium, filtered, GFAA, labcode 1936, parmcode 1030I |
| From QADATA Summary November 2004 | A positive bias is still evident (9.5%). Since late-June, 35 out of 36 values (excluding one outlier) are above the MPV. A bias of this magnitude is not evident for chromium, WWR by GFAA. | NWQL QC charts indicate a <1% negative bias. New standards were prepared and verified on September 27th, 2004. A NIST traceable standard was run on the GFAA. It indicates that the new standards are within specification. Data indicates the bias is < 2%. Six of the blind samples were rerun on November 4th, and a positive bias is not evident. | NWQL will continue to closely monitor analysis for a bias. |
| From QADATA Summary September 2004 | The chart shows a positive bias (9.3%) beginning June 23, 2004. Since this time, all results (19 values) are higher than the MPV. | NWQL QC charts do not indicate a bias. Eleven of the BQS samples were run on the ICP and all values were higher than the MPV. The QC for the run (CCVs and TPCs) was running slightly lower than the MPV. A NIST traceable standard was run on the GFAA. The values came out slightly lower than the MPV. | NWQL will continue to closely monitor analysis. |
| Chromium, Filtered ICP labcode 0722 parmcode 1030E | Data exhibit an upward trend starting November 2004. | We have reviewed internal QC data and we do not have an explanation for the bias observed by BQS. NWQL QC charts show a very slight bias, but within acceptance criteria. | We have no explanation for the bias. We will continue to monitor the situation and attempt to find the cause and review the data with BQS in a month. |
| Copper, Filtered GFAA labcode 1558 parmcode 1040F | Due to the low number of requests for this analytical test, it will no longer be monitored by IBSP. |  |  |
| Previous comments for copper, filtered, GFAA, labcode 1558, parmcode 1040F |
| From QADATA Summary November 2004 | As of mid-May 2004, the data show an increase in variability. Three out of 21 values (14%) are outside of +/-2 f-pseudosigma. In addition, data in that time range show a positive bias (10.0%). | NWQL QC charts also show an increase in variability during June and July, although all QC data was within acceptable limits. Variability has improved since then. A positive bias is evident on QC charts of 2-4% from May to September. New standards were prepared and verified on September 27th, 2004 and the current bias is <2%. | NWQL will continue to closely monitor analysis for bias and variability. |
| Copper, Whole-Water Recoverable GFAA labcode 1559 parmcode 1042F | A positive bias is still demonstrated. Forty of the last 43 results (over the last 10 months) are higher than their target value. | NWQL QC charts continue to show slight positive bias, but within acceptance criteria since standards were prepared September 27, 2004. | We will work with BQS to resolve the bias issue. |
| Previous comments for copper, whole-water recoverable, GFAA, labcode 1559, parmcode 1042F |
| From QADATA Summary November 2004 | Data show a positive bias (12.8%) throughout the last six months. All results (25 values) are greater than the MPV. | NWQL QC charts from May to September show a positive bias of about 5%. New standards were prepared and verified on September 27th, 2004. The current bias is <1%. | NWQL will continue to closely monitor analysis for a bias. |
| Copper, Whole-Water Recoverable ICP-MS labcode 2379 parmcode 1042H | Data exhibit an upward trend starting mid-November 2004. Analysis should be closely monitored until the next QADATA Summary is distributed. | The NWQL QC charts do not show an upward trend. | The NWQL will review and compare internal QC information with BQS in one month. The NWQL will work with BQS if the trend continues. |
| Manganese, Filtered ICP-MS labcode 1793 parmcode 1056G | Data still demonstrate a negative bias (7.3%). Twenty-nine of the last 33 results (since September 2004) are lower than their target value. This bias is shown also for the whole-water recoverable method by ICP-MS but is not evident for ICP. | The standards made on November 15th were verified and are within acceptable limits. Internal QC charts are within acceptable limits but have slightly low bias for the 25 ug/L QC standard and slightly high bias for the 250 ug/L QC standard. | The NWQL will work with BQS and Methods Research and Development to determine the cause of the bias observed on the BQS and NWQL charts. |
| Previous comments for manganese, filtered, ICP-MS, labcode 1793, parmcode 1056G |
| From QADATA Summary November 2004 | Data since late-August 2004 indicate a negative bias (7.7%). Since this time, 12 out of 13 results are lower than the MPV. | NWQL QC charts indicate a negative bias of about 3.5% since new standards were prepared and verified on September 17th, 2004. The working standards are made every two months. The stock does not appear to have a bias. | New working standards will be made the week of November 15, 2004 and as long as the bias is less than <2%, they will be put into use, otherwise they will be remade again. NWQL will continue to closely monitor analysis for a bias. |
| Manganese, Whole-Water Recoverable ICP-MS labcode 2382 parmcode 1055C | Data indicate a negative bias (8.4%). All results since July 2004 are lower than their target value and the bias seems to be getting more pronounced. This bias is shown also for the filtered method by ICP-MS but is not evident for ICP. | The internal QC chart confirms the bias in the BQS chart. | The NWQL will work with BQS and Methods Research and Development to determine the cause of the bias observed on the BQS and NWQL charts. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | The variability has increased to where almost 50% of the results received since the last QADATA Summary is outside of +/-2 f-pseudosigma. | The NWQL QC charts continue to show the same variability observed in the November QADATA Summary; however, the 3.2% bias has decreased. The QC for the runs with the outlier BQS samples was acceptable. There were no carryover problems or sample mispours. | We have no explanation for the high variability observed in the BQS charts. We will continue to monitor for variability and review with BQS in one month. |
| Previous comments for strontium, filtered, ICP-MS, labcode 2507, parmcode 1080I |
| From QADATA Summary November 2004 | Results still show high variability. Since mid-June 2004, 14 out of 45 values (31%) are outside of –2 f-pseudosigma. A negative bias is still demonstrated (5.1%). | NWQL QC charts indicate a positive bias of about 3.2%. Using a NIST standard as a check, the calibration standards are also showing a positive bias. The section is looking into the possibility that the concentration for strontium is not correctly calculated in the custom mix that is purchased from an outside vendor. NWQL QC charts indicate a decrease in variability beginning in late September, 2004. We do not appear to have high variability at this time. | New standards will be purchased this month and verified by a NIST third party check standard. In the interim, we will be using an individual strontium standard to make an interim calibration standard. NWQL will continue to closely monitor analysis for bias and variability. |
| From QADATA Summary September 2004 | Results received as of mid-June 2004 indicate increased variability. Six out of 27 values (22%) are outside of –2 f-pseudosigma. The data also exhibit a slight negative bias (4.3%). | A software problem was detected in mid-August that caused a slight increase in variability for strontium. Also, at this same time, it was determined that the instrument detector needed replacement. NWQL QC charts indicate a negative bias of 5%. | The instrument vendor has fixed the software problem and a new detector has been installed. Calibration standards will be remade and verified by September 14th. NWQL will continue to monitor strontium for bias and variability. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Chloride, Filtered IC labcode 1571 parmcode 0940J | Results received since December 2004 exhibit no bias and good variability. No corrective action necessary. |  |  |
| Previous comments for chloride, filtered, IC, labcode 1571, parmcode 0940J |
| From QADATA Summary November 2004 | The data indicate a very slight positive bias starting around June 2004 and increasing again in September. The overall bias since September is 3.2%, and this period is marked by an increase in variability (5 out of the last 18 values are greater than +2 f-pseudosigma). | We reviewed the run QC and we do not see the same trends that are evident in the BQS samples. | We are working with BQS to identify the samples and we will investigate each one to identify any anomalies. |
| Magnesium, Filtered ICP labcode 0663 parmcode 0925C | Results received since November 2004 exhibit no bias and good variability. No corrective action necessary. |  |  |
| Previous comments for magnesium, filtered, ICP, labcode 0663, parmcode 0925C |
| From QADATA Summary November 2004 | The negative bias (5.1%) starting mid-May 2004, still persists. Since this time there have been 67 out of 69 results below the MPV. | The negative bias is not evident on NWQL QC charts at the 0.5 mg/L level, which is the concentration where our TPC and CCV samples are at. However, a slight negative bias is evident at the 5 mg/L using a NIST standard as a check. New standards were prepared on October 29th, 2004. Current QC data indicate the bias is <2%. | NWQL will continue to closely monitor magnesium for a bias. |
| Potassium, Filtered ICP labcode 2773 parmcode 0935C | A positive bias is shown starting November 2004 (5.7%). As of November, 30 out of 31 results are higher than their target value. | The NWQL QC charts show less bias since January 2005. The charts also show a decrease in variability since October 2004. | New standards will be prepared and verification initiated by Feb. 14, 2005. We will continue to monitor for bias after the new standards are prepared and review the data with BQS in a month. |
| Silica, Filtered ICP labcode 0667 parmcode 0955D | The previously noted bias has decreased to less than 3% as of November 2004. No corrective action necessary. |  |  |
| Previous comments for silica, filtered, ICP, labcode 0667, parmcode 0955D |
| From QADATA Summary November 2004 | The data indicate a very slight positive bias starting mid-June 2004 and increasing again in September. The overall bias since September is 7%. | A slight positive bias of <2% was evident on NWQL QC charts as well. New standards were prepared on October 23rd, 2004 in response to a steady increase in TPC and CCV values. The bias is <1% on in-house QC charts. | NWQL will continue to closely monitor analysis. |
| Sulfate, Filtered IC labcode 1572 parmcode 0945G | The data show the negative bias returning. Results received since November 2004 are biased negative by 7.2%. | We are aware of the slight negative bias. Bench QC has shown negative bias at concentrations less than about 10 mg/L. | We are currently trouble shooting all aspects of the analytical method in an effort to identify the source of the bias. As we make adjustments, we will continue to monitor for improvements in the bias. |
| Previous comments for sulfate, filtered, IC, labcode 1572, parmcode 0945G |
| From QADATA Summary November 2004 | Data as of mid-August 2004 exhibit no bias and excellent variability. No corrective action necessary. |  |  |
| From QADATA Summary September 2004 | The chart indicates a slight negative bias (4.7%) starting early-April 2004. All 29 results since this time (except one outlier) are below the MPV. | We reviewed internal QC to help evaluate the observed bias. Our data indicates a bias of about 3% or less, in only the midsection of the analytical range. We also noticed increased variability during the period of the observed bias. We are currently implementing new instruments on this analytical line, and will eliminate use of the instrument producing the observed bias in the next few months. | We have no explanation of the bias, or why it is only evident in the middle of the analytical range. We will continue to monitor the situation and attempt to find the cause. |