##### QADATA Summary – April 2005

The Inorganic Blind Sample Project (IBSP) reviews about 110 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 09/01/04 through 04/07/05; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/FY05charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at  http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Chromium, Filtered, ICP, Labcode 0722, Parmcode 1030E	4

Copper, Whole-Water Recoverable, GFAA, Labcode 1559, Parmcode 1042F	5

Copper, Whole-Water Recoverable, ICP-MS, Labcode 2379, Parmcode 1042H	5

Lithium, Filtered, ICP, Labcode 0664, Parmcode 1130B	5

Manganese, Filtered, ICP-MS, Labcode 1793, Parmcode 1056G	6

Manganese, Whole-Water Recoverable, ICP-MS, Labcode 2382, Parmcode 1055C	6

Strontium, Filtered, ICP-MS, Labcode 2507, Parmcode 1080I	7

Mercury, Whole-Water Recoverable, CVAF, Labcode 2708, Parmcode 71900D	8

Zinc, Whole-Water Recoverable, ICP-MS, Labcode 2390, Parmcode 1092D	8

Fluoride, Filtered, ISE, Labcode 0031, Parmcode 0950B	8

Potassium, Filtered, ICP, Labcode 2773, Parmcode 0935C	9

Sulfate, Filtered, IC, Labcode 1572, Parmcode 0945G	9

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Chromium by ICP, filtered, parmcode 1030E – Positive bias starting November 2004.

- Copper by GFAA, whole-water recoverable, parmcode 1042F – Long-term positive bias starting around May 2004.

- Copper by ICP-MS, whole-water recoverable, parmcode 1042H – Positive bias starting around mid-November 2004.

- Fluoride by ISE, filtered, parmcode 0950B – Positive bias starting mid-November 2004.

- Lithium by ICP, filtered, parmcode 1130B – Positive bias starting December 2004.

- Manganese by ICP-MS, whole-water recoverable, parmcode 1055C– Previous negative bias (8%) is now about 12%.

- Zinc by ICP-MS, whole-water recoverable, parmcode 1092D – Negative bias starting December 2004.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Chromium, Filtered ICP labcode 0722 parmcode 1030E | Data still demonstrate a positive bias that started November 2004. All results since then (20 values) are greater than the target value, and most results since December are greater than +1 f-pseudosigma. (Note: Chromium by GFAA also demonstrates a positive bias; however, most values are less than +1 f-pseudosigma.) | Standards were prepared and verified in February. The slight bias on internal control charts has decreased starting about March 2005. Upon request by BQS, samples analyzed before the bias appeared were re-analyzed after the new standards were prepared. The results duplicate within the precision of the method, suggesting the new standards may have corrected the bias. | We believe the bias has been corrected and will continue to monitor. |
| Previous comments for chromium, filtered, ICP, labcode 0722, parmcode 1030E |
| From QADATA Summary February 2005 | Data exhibit an upward trend starting November 2004. | We have reviewed internal QC data and we do not have an explanation for the bias observed by BQS. NWQL QC charts show a very slight bias, but within acceptance criteria. | We have no explanation for the bias. We will continue to monitor the situation and attempt to find the cause and review the data with BQS in a month. |
| Copper, Whole-Water Recoverable GFAA labcode 1559 parmcode 1042F | Results received since the last QADATA Summary suggest the positive bias is still increasing. Fifty-two of the last 55 results (over the last 12 months) are higher than the target value. In addition, variability seems to be increasing. Four out of the last 19 results (21%) are outside +/-2 f-pseudosigma. | The slight bias on internal QC charts described in the last summary has lessened since mid-January 2005. The in-bottle digestion spike charts show no bias or changes in variability from October 2004 to present. The QC for the runs with the outlier BQS samples was acceptable. There were no carryover problems or sample mispours. | We have no explanation for the continued bias and variability, and will continue to work with BQS to resolve the bias and variability. |
| Previous comments for copper, whole-water recoverable, GFAA, labcode 1559, parmcode 1042F |
| From QADATA Summary February 2005 | A positive bias is still demonstrated. Forty of the last 43 results (over the last 10 months) are higher than the target value. | NWQL QC charts continue to show slight positive bias, but within acceptance criteria since standards were prepared September 27, 2004. | We will work with BQS to resolve the bias issue. |
| From QADATA Summary November 2004 | Data show a positive bias (12.8%) throughout the last six months. All results (25 values) are greater than the MPV. | NWQL QC charts from May to September show a positive bias of about 5%. New standards were prepared and verified on September 27th, 2004. The current bias is <1%. | NWQL will continue to closely monitor analysis for a bias. |
| Copper, Whole-Water Recoverable ICP-MS labcode 2379 parmcode 1042H | Results received since the last QADATA Summary show the positive bias still present. As of mid-November 2004, data are biased 23% high and all values except one are greater than the target value. | As discussed at a meeting with BQS, the internal charts have shown no bias and no changes in variability in performance going back to October 2003. | We have no explanation for the continued bias and suggest further work with BQS to ascertain the cause of bias. |
| Previous comments for copper, whole-water recoverable, ICP-MS, labcode 2379, parmcode 1042H |
| From QADATA Summary February 2005 | Data exhibit an upward trend starting mid-November 2004. Analysis should be closely monitored until the next QADATA Summary is distributed. | The NWQL QC charts do not show an upward trend. | The NWQL will review and compare internal QC information with BQS in one month. The NWQL will work with BQS if the trend continues. |
| Lithium, Filtered ICP labcode 0664 parmcode 1130B | Results as of December 2004 are biased 12% high. The most recent data show no bias; however, these values are less than the Laboratory Reporting Limit. Analysis should be closely monitored until the next QADATA Summary is distributed. | Upon request by BQS, samples analyzed before November 2004 were re-analyzed in April. The results duplicate within the precision of the method. | We will continue to monitor and review with BQS in one month. |
| Manganese, Filtered ICP-MS labcode 1793 parmcode 1056G | The previously noted bias has improved. As of January 2005, the data are biased by less than 3%. Analysis should be closely monitored until the next QADATA Summary is distributed. | The calibration standards prepared in March 2005 from single elements have replaced the multi-element standards used previously. This appears to have improved the bias. | We believe the bias has been corrected and will continue to monitor. |
| Previous comments for manganese, filtered, ICP-MS, labcode 1793, parmcode 1056G |
| From QADATA Summary February 2005 | Data still demonstrate a negative bias (7.3%). Twenty-nine of the last 33 results (since September 2004) are lower than the target value. This bias is shown also for the whole-water recoverable method by ICP-MS but is not evident for ICP. | The standards made on November 15th, 2004 were verified and are within acceptable limits. Internal QC charts are within acceptable limits but have slightly low bias for the 25 ug/L QC standard and slightly high bias for the 250 ug/L QC standard. | The NWQL will work with BQS and Methods Research and Development to determine the cause of the bias observed on the BQS and NWQL charts. |
| From QADATA Summary November 2004 | Data since late-August 2004 indicate a negative bias (7.7%). Since this time, 12 out of 13 results are lower than the MPV. | NWQL QC charts indicate a negative bias of about 3.5% since new standards were prepared and verified on September 17th, 2004. The working standards are made every two months. The stock does not appear to have a bias. | New working standards will be made the week of November 15, 2004 and as long as the bias is less than <2%, they will be put into use, otherwise they will be remade again. NWQL will continue to closely monitor analysis for a bias. |
| Manganese, Whole-Water Recoverable ICP-MS labcode 2382 parmcode 1055C | All results since March 2004 (35 values) are lower than the target value, and the negative bias seems to be increasing in magnitude. Results as of December 2004 are biased low by 12%. A bias of this magnitude is not evident for Mn by ICP. | The bias still exists, although not as large as observed on the BQS charts. New single element calibration standards will be prepared. If the new standards are acceptable, we believe the bias will improve as it has for the manganese, filtered charts. Discussions with Methods Research and Development did not result in any additional suggestions for corrective action. | We will prepare and verify new calibration standards by May 6, 2005 and will continue to monitor. |
| Previous comments for manganese, whole-water recoverable, ICP-MS, labcode 2382, parmcode 1055C |
| From QADATA Summary February 2005 | Data indicate a negative bias (8.4%). All results since July 2004 are lower than their target value and the bias seems to be getting more pronounced. This bias is shown also for the filtered method by ICP-MS but is not evident for ICP. | The internal QC chart confirms the bias in the BQS chart. | The NWQL will work with BQS and Methods Research and Development to determine the cause of the bias observed on the BQS and NWQL charts. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | Results received since the last QADATA Summary show an improvement in variability. Since the last summary only one result out of 25 is outside +/-2 f-pseudosigma. Analysis should be closely monitored until the next QADATA Summary is distributed. | The variability has improved since the last QA data summary. | We will continue to monitor and review in one month with BQS. |
| Previous comments for strontium, filtered, ICP-MS, labcode 2507, parmcode 1080I |
| From QADATA Summary February 2005 | The variability has increased to where almost 50% of the results received since the last QADATA Summary is outside of +/-2 f-pseudosigma. | The NWQL QC charts continue to show the same bias (about 3.2%) observed in the November QA Data summary. The QC for the runs with the outlier BQS samples was acceptable. There were no carryover problems or sample mispours. | We have no explanation for the high variability observed in the BQS charts. We will continue to monitor for variability and review with BQS in one month. |
| From QADATA Summary November 2004 | Results still show high variability. Since mid-June 2004, 14 out of 45 values (31%) are outside of –2 f-pseudosigma. A negative bias is still demonstrated (5.1%). | NWQL QC charts indicate a positive bias of about 3.2%. Using a NIST standard as a check, the calibration standards are also showing a positive bias. The section is looking into the possibility that the concentration for strontium is not correctly calculated in the custom mix that is purchased from an outside vendor. NWQL QC charts indicate a decrease in variability beginning in late September, 2004. We do not appear to have high variability at this time. | New standards will be purchased this month and verified by a NIST third party check standard. In the interim, we will be using an individual strontium standard to make an interim calibration standard. NWQL will continue to closely monitor analysis for bias and variability. |
| From QADATA Summary September 2004 | Results received as of mid-June 2004 indicate increased variability. Six out of 27 values (22%) are outside of –2 f-pseudosigma. The data also exhibit a slight negative bias (4.3%). | A software problem was detected in mid-August that caused a slight increase in variability for strontium. Also, at this same time, it was determined that the instrument detector needed replacement. NWQL QC charts indicate a negative bias of 5%. | The instrument vendor has fixed the software problem and a new detector has been installed. Calibration standards will be remade and verified by September 14th. NWQL will continue to monitor strontium for bias and variability. |
| Mercury, Whole-Water Recoverable CVAF labcode 2708 parmcode 71900D | Data for March 2005 are biased 13% low. The most recently received data show no bias. Analysis should be closely monitored until the next QADATA Summary is distributed. | The internal charts confirm a bias that stopped around the beginning of April 2005. | We will continue to monitor and review with BQS in one month. |
| Zinc, Whole-Water Recoverable ICP-MS labcode 2390 parmcode 1092D | Data since December 2004 are biased negative by 6.3%. All results as of then (21 values) are lower than the target value. This bias is not evident for filtered zinc by ICP-MS. | One internal chart shows a -4% bias for December 2004 only and no bias is observed on all charts for the remaining time period. | We have no explanation for the low bias observed in the BQS charts. We will continue to monitor for the low bias and review with BQS in one month. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Fluoride, Filtered ISE labcode 0031 parmcode 0950B | Data since mid-November 2004 are biased high by 6.2%. Forty-three out of the last 44 results are greater than the target value. | The NWQL agrees with the observation. | We are ordering new stock standards and will prepare and verify new working standards by April 29, 2005. We also ordered a new electrode to see if the current electrode is contributing to the bias. |
| Potassium, Filtered ICP labcode 2773 parmcode 0935C | The previously noted bias has decreased to less than 3% as of February 2005. No corrective action necessary. | The standards prepared in February appear to have corrected the bias observed in the BQS chart. | We will continue to monitor for bias and review the data with BQS in a month. |
| Previous comments for potassium, filtered, ICP, labcode 2773, parmcode 0935C |
| From QADATA Summary February 2005 | A positive bias is shown starting November 2004 (5.7%). As of November, 30 out of 31 results are higher than their target value. | The NWQL QC charts show less bias since January 2005. The charts also show a decrease in variability since October 2004. | New standards will be prepared and verification initiated by Feb. 14, 2005. We will continue to monitor for bias after the new standards are prepared and review the data with BQS in a month. |
| Sulfate, Filtered IC labcode 1572 parmcode 0945G | The previously noted bias is no longer evident as of February 2005. No corrective action necessary. |  |  |
| Previous comments for sulfate, filtered, IC, labcode 1572, parmcode 0945G |
| From QADATA Summary February 2005 | The data show the negative bias returning. Results received since November 2004 are biased negative by 7.2%. | We are aware of the slight negative bias. Bench QC has shown negative bias at concentrations less than about 10 mg/L. | We are currently trouble shooting all aspects of the analytical method in an effort to identify the source of the bias. As we make adjustments, we will continue to monitor for improvements in the bias. |
| From QADATA Summary November 2004 | Data as of mid-August 2004 exhibit no bias and excellent variability. No corrective action necessary. |  |  |
| From QADATA Summary September 2004 | The chart indicates a slight negative bias (4.7%) starting early-April 2004. All 29 results since this time (except one outlier) are below the MPV. | We reviewed internal QC to help evaluate the observed bias. Our data indicates a bias of about 3% or less, in only the midsection of the analytical range. We also noticed increased variability during the period of the observed bias. We are currently implementing new instruments on this analytical line, and will eliminate use of the instrument producing the observed bias in the next few months. | We have no explanation of the bias, or why it is only evident in the middle of the analytical range. We will continue to monitor the situation and attempt to find the cause. |