##### QADATA Summary – July 2003

The Inorganic Blind Sample Project (IBSP) reviews 95 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 84 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-assurance samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-assurance charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 11/15/02 through 06/20/03; however, longer-range charts are shown when they add clarity to the problem.  The attached quality-assurance charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality control-review as the data released to the districts.  Thus, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-assurance charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have deteriorated or are known to have been affected by a problem at IBSP are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report to help provide a clearer perspective of the overall concern.

Charts containing all data received as of 07/07/03 may be accessed at http://bqs.usgs.gov/bsp/WY03charts.htm.  Lab personnel and customers can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp/mainpage.html (select QADATA Application button).  Charts pertaining specifically to this report can be accessed by following the blue, underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Ammonia, Filtered, Colorimetric, Labcode 1976,  Parmcode 0608F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	3

Orthophosphate, Filtered, Colorimetric, Labcode 1974,  Parmcode 0671H . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Barium, Whole-Water Recoverable, ICP, Labcode 2352, Parmcode 1007C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .   	5

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .   5

Lead, Filtered, GFAA, Labcode 1560, Parmcode 1049F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

Lead, Whole Water Recoverable, GFAA, Labcode 1561, Parmcode 1051F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	6

Strontium, Filtered, ICP, Labcode 0652, Parmcode 1080B . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

Calcium, Filtered, ICP, Labcode 0659, Parmcode 0915D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	7

OWQRL

Analyte	Page

Arsenic, Whole Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Barium, Filtered, ICP, Labcode 3002, Parmcode 1005(4) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	8

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Ammonia, Filtered Colorimetric labcode 1976 parmcode 0608F | The low bias demonstrated during March 2003 has progressively improved throughout April. No corrective action necessary. | None needed. | The NWQL will continue to monitor the IBSP control charts for this parameter. |
| Previous comments for ammonia, filtered, labcode 1976, parmcode 0608F |
| From QADATA Summary May 2003 | A negative bias has developed that started in March 2003. This bias is not shown by the OWQRL. | The low bias evident in March and early April appears to be improving and the mid-April data are close to the median. | The Nutrients Unit is investigating the cause of the bias. A new reference sample will be prepared that is near the concentration of NS10 for on-line monitoring. The NWQL will continue to monitor the on-line and IBSP charts for this parameter and look for corrective actions if indicated. |
| Orthophosphate, Filtered Colorimetric labcode 1974 parmcode 0671H | Over the last two months (April and May 2003) results for NS10 show improvement relative to earlier data. Analysis should be closely monitored to ensure values stay within +/- 2 f-pseudosigma. | Values for N77 and N78 are distributed closely around the median. QC samples prepared from the NIST certified PO4 reference material purchased by the NWQL also run close to their certified value. | The Nutrients Unit will continue to run NIST samples and monitor the IBSP charts. |
| Previous comments for orthophosphate, filtered, labcode 1974, parmcode 0671H |
| From QADATA Summary May 2003 | N76 has been replaced by a new matrix (N78) which now indicates that this line has improved variability. However, for the past two months, about 50% of the data for NS10 are outside of -2 f-pseudosigma, whereas for the OWQRL no data for NS10 are outside of 2 f-pseudosigma. | The NIST certified reference sample purchased by the NWQL was prepared at concentrations close to the PO4 concentration in N75, N76 and NS10. The NIST sample near the NS10 level ran slightly high on average and the NIST sample near the N75 level gave results that averaged close to the certified value. Results for the NIST sample near the N-76 concentration indicate a higher than expected concentration. All NIST samples were cross-checked on the low-level ortho-P line with results that were close to the regular line values. | The Nutrients Unit will continue to analyze NIST samples in addition to the on-line QC samples. The results from these samples and the IBSP results will continue to be monitored. |

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Barium, Whole-Water Recoverable ICP labcode 2352 parmcode 1007C | The historical negative bias has not improved and may be getting worse. It is again recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | The NWQL agrees with IBSP that the barium bias is getting worse. New ICP instrumentation has come online for the dissolved line. | The NWQL will continue to monitor the barium WWR line on the different instrument to see if there is an improvement. The ICP WWR method has been moved to a different instrument that has had recent maintenance. We will reanalyze IBSP samples on the this instrumentation. We plan to have the new Perkin Elmer ICP instrumentation online this summer. |
| Previous comments for barium, whole-water recoverable, ICP, labcode 2352, parmcode 1007C |
| From QADATA Summary May 2003 | The historical negative bias persists. It is recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | Online QC charts show that the negative bias increased significantly during mid-March. Several troubleshooting measures were checked on the instrument including: plasma power, pump efficiency, cleaning of nebulizer and spray chamber, dark current test, all of which showed no indication of a problem. The only test that showed a significant problem was that the Purged Optical Path, which needed adjusting. | The NWQL will reanalyze samples on the new ICPs to compare the data, once we have assessed the WWR prove-out data for the new instruments. We will continue to monitor online QC charts to see if the negative bias has improved as a result of the optical path adjustment. |
| From QADATA Summary March 2003 | Although there are two points above the median, this line is still exhibiting a negative bias. No new values have been received since the implementation of new standards (last week of February 2003). Analysis should be closely monitored to determine the effects of the new standards. | New standards were prepared and verified by ICPMS. New standards were accepted. | The NWQL will continue to monitor. More data from the IBSP since these actions were implemented will be examined in order to better assess if the bias is improving. |
| From QADATA Summary January 2003 | The historical negative bias persists. | NWQL charts still show a slight negative bias. The standards approved in August, 2002 were recently re-checked and found to be slightly outside acceptance limits on a new ICP instrument in development status. | New standards will be prepared and checked the first week in February, 2003. They will be compared with the old standards on the ICP as well as by ICPMS (alternate method). SRS performance will be evaluated using the new standards. The OWQRL lab will be called to compare procedures as well. |
| From QADATA Summary September 2002 | Magnitude of historical negative bias has decreased although a negative bias is still demonstrated. Analysis should still be closely monitored to ensure magnitude of bias does not increase. | Adjustments to past descriptions of bias from IBSP have been to digest the calibration standards and QC to match the digestion process of the samples. | Calibration standards were made and accepted on August 20, 2002. Standards will be cross-checked by a different method. Using lab QC data and IBSP, we will continue to monitor and, additionally, obtain more current SRWS from BQS for internal lab QC. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | As of January 2003, a high bias has developed and variability increased. | NWQL QC charts show variable but acceptable data. | Instrument has recently been optimized, and the NWQL will continue to monitor. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | The positive bias, although still present, has improved. All data received since April 2003 are equal to or within one f-pseudosigma of the median. Analysis should be closely monitored until next QADATA Summary is distributed. | Bias may have been result of old standards | New standards are being prepared and compared to purchased NIST standards. The line will be monitored for any change. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary May 2003 | The positive bias, although showing improvement, is still present. | Since new calibration standards were made in January, third party QC charts show a reduction in bias from 5 percent to 1 percent at 2 ug/L and 3 percent to 1 percent at 20 ug/L. Blank control charts show no bias. The SRWS T-165, T-167, and T-169 control charts still show larger relative positive bias. It is possible the purchased cadmium standards in current use are of lesser quality—a problem not unlike what we have recently encountered with purchased standards used for antimony. | Different sources of standards will be obtained and tested in attempt to reduce bias. NWQL will continue monitoring with older SRWS that have been historically less biased for cadmium along with the newer SRWS where a larger bias is observed. |
| From QADATA Summary March 2003 | Positive bias still evident although the two most recently received values are close to the median. Cadmium by other methods at the NWQL and the OWQRL do not show the significant bias that this line does. | Third party QC charts show improvement since new standards started in January 2003. | The lab will continue to closely monitor third party and other quality assurance samples. |
| From QADATA Summary January 2003 | The slight positive bias that developed around April 2002 appears to have increased even more around mid-September 2002. | Third party QC charts show a 5% bias at 2 ug/L and 3% bias at 20 ug/L since October 2002. All in-house QC data have been acceptable. New SRSs were started December 2002. Not enough data for the newest SRSs to evaluate bias. | New standards are acceptable and have been in use since January 2003. The NWQL will monitor cadmium closely for improvement in bias. |
| From QADATA Summary November 2002 | A slight positive bias that developed around mid-April 2002 is increasing. | In-house charts show bias starting at about the same time. All in-house QC data has been acceptable. Bias most noticeable in the two newest mixes. | Newly purchased standards are prepared and currently undergoing verification against the newest SRS and other third party standards. The lab will monitor cadmium closely for improvement. |
| Lead, Filtered GFAA labcode 1560 parmcode 1049F | The most recently received data suggest that the variability is not improving and may be increasing. | Variable, but acceptable data quality is also noted on NWQL charts. | Instrument has recently been optimized and the NWQL will monitor for improvement. |
| Previous comments for lead, filtered, GFAA, labcode 1560, parmcode 1049F |
| From QADATA Summary May 2003 | Variability has increased somewhat since January 2003. Variability should be closely monitored until next QADATA Summary is distributed. | Change in variability not noticeable on NWQL charts. | NWQL will continue to monitor. |
| Lead, Whole Water Recoverable GFAA labcode 1561 parmcode 1051F | The data indicate increased variability and positive bias starting March 2003. | Variable, but acceptable data quality is also noted on NWQL charts. | Instrument has recently been optimized and the NWQL will monitor for improvement. |
| Strontium, Filtered ICP labcode 0652 parmcode 1080B | The variability has increased. Analysis should be closely monitored until the next QADATA Summary is distributed. | The NWQL has changed to new ICP instrumentation. | QC will be monitored on the new instrumentation until the next QADATA Summary is distributed. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered ICP labcode 0659 parmcode 0915D | The data show increased variability with a slight overall negative bias (5%) from December 2002 through mid-April. All data received as of mid-April indicate no problems. Analysis should be closely monitored until next QADATA Summary is distributed. | The NWQL has changed to new ICP instrumentation. | QC will be monitored on the new instrumentation until the next QADATA Summary is distributed. |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole Water Recoverable GFAA labcode 3119 parmcode 1002(1) | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode 1002(1) |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 SD with 3 of 13 (23%) outside of +/-3 SD. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 SD. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |
| Barium, Filtered ICP labcode 3002 parmcode 1005(4) | The historical positive bias appears to be increasing. | The positive bias could be caused by a change in the background around the barium emission line. | A change in the point at which background correction is applied is being tested. |