ARCADIS G&M, Inc.

4915 Prospectus Drive

Suite F

Durham

North Carolina 27713

Tel 919 544 4535

Fax 919 544 5690

MEMO

| To: Victor D’Amato | Copies: Copies |
| --- | --- |
| From: |  |
| Kim Egler |  |
| Date: | ARCADIS Project No.: |
| 21 June 2004 | RN990230.0005.00001 |
| Subject: |  |
| Validation of Wipe Samples from CCA Treated Wood Collected April 13, 2004 |

Two hundred and ninety-nine (299) samples were collected on April 13, 2004 from the CAA Treated Wood Minidecks located at the USEPA in Research Triangle Park, North Carolina.  These samples were triple-extracted with 10 percent nitric acid by ARCADIS personnel and submitted by overnight courier at ambient temperature on April 14, 2004 to STL Savannah, Savannah, Georgia.  Samples were analyzed on May 7-8 and 11-13, 2004.  This validation covers the samples submitted for arsenic (As), chromium (Cr), and copper (Cu) analysis by USEPA Method 200.8 (ICP-MS, modified).  This methodology was modified to utilize hydrogen plasma, rather than argon, as classically performed.  This modification eliminates concerns over the formation of Ar40Cl35, which can create a positive bias when measuring As.

Sample Delivery Group (SDG) WIPE10, Project Number S443176

SS-1315 through SS-1403

SS-1405 through SS-1614 (sample SS-1432 in this sequence is actually SS-1432A)
SS-1432B

Validation of this data was performed following the quality assurance/quality control (QA/QC) criteria set forth in the “USEPA Contract Laboratory Program (CLP) National Functional Guidelines (NFG) for Inorganic Data Review,” July 2002.  The NFG have slightly different criteria for some parameters than does Method 200.8.  In these cases, the more stringent criteria were used.

# Metals Inorganic Review

# I.	Preservation and Holding Times

The samples arrived at the laboratory promptly and were analyzed within the analytical holding time.  The samples had been shipped at ambient temperature after being triple-extracted in 10 percent nitric acid.  Although the NFG recommends maintaining aqueous samples at 4ºC±2ºC, no action is specified for this condition, and no action was taken.  The Case Narrative included sample identifications and indicated that all samples were analyzed at various dilutions for As, Cr, and Cu, due to the abundance of these metals.

# II.	ICP-MS Tune Analysis

As required, 100 ug/L (ppb) tune checks were run on the required isotopes before each calibration.  Each tune was run five times in a row, exceeding the four times in a row criteria required by the method.  The percent Relative Standard Deviation (%RSD) was less than 5% for all isotopes, and the mass calibrations and peak widths met control criteria.

# III.	Calibration

The instrument was successfully calibrated before these samples were analyzed.  The two, four-point calibrations ranged from 0.5 to 100 ug/L.

Initial Calibration Verification (ICV) samples were run immediately following each calibration, and all three metal recoveries fell within the 90-110% control limits.  Continuing Calibration Verification (CCV) samples were run at a frequency of 10 percent, and at the beginning of the run and after the last sample, as required.  All recoveries were within established control limits.

The Contract Required Quantitation Limit (CRQL) sample, referred to as the CRDL in the lab report, was run as the first sample in the run, immediately preceding the Interference Check Samples (ICS), as stipulated in the NFG.  Although the NFG states that the CRQL should be run at the beginning and end of each sample analysis run and every 20 analytical samples, the method has no such requirement.  The recoveries for As, Cr, and Cu were within control limits.

# IV.	Blanks

An Initial Calibration Blank (ICB) was run following each ICV, and was nondetect for any of the target metals.  Continuing Calibration Blanks (CCB) were run immediately following each CCV at a frequency of 10 percent, as well as at the beginning of the run and after the last CCV.  All of the CCBs were nondetect for the three metals, with the following exceptions.

Samples analyzed May 7-8. 2004:

CCB	Metal Detected (ug/L)	Action Level (ug/L)

CCB#12	As – 0.106	0.530

CCB#17	As – 0.151	0.755

CCB#19	As – 0.106	0.530

CCB#20	As – 0.100	0.500

Samples analyzed May 11-13. 2004:

CCB	Metal Detected (ug/L)	Action Level (ug/L)

CCB#3	As – 0.153	0.765

However, all associated sample results were either nondetect for As or were above the Action Level, and no action was required.

Because the samples were extracted prior to shipping to the laboratory, there were no Preparation (Method) Blanks associated with these samples, and the CCBs were used as the Method Blanks.  Fifteen field blanks (extraction fluid only) were included in this SDG.  All but one of the blanks detected Cu and/or Cr, and resulted in the following samples – which fell below the associated Action Level – being qualified as nondetect:

Cr	SS-1422, SS-1427, SS-1432A, SS-1448, SS-1450, and SS-1459

In addition, one 10 percent nitric blank was run (SS-1586) and was nondetect for all three target metals.

# V.	Inductively Coupled Plasma-Interference Check Sample (ICS)

Interference Check Samples (ICS) were run at the beginning of each of the three runs, after the ICV, and were followed by a CCV and CCB, as required.  The recoveries fell within the control limits and gave no indication of isobaric interferences.

# VI.	Laboratory Control Sample (LCS)

Because the laboratory did not extract the samples, the CCVs were used as the LCS samples.

# VII.	Duplicate Sample Analysis

Duplicate samples – like LCS samples – are to be carried through the entire preparation and analysis process. Because the laboratory did not extract the samples, no duplicate analyses were performed.

# VIII.	Spike Sample Analysis

Because the laboratory did not extract the samples, Matrix Spike (MS)/MS Duplicate (MSD) samples could not be analyzed.  Therefore, Post Digestion Spike (PDS) samples were run at a frequency of 10 percent.  All recoveries were within control limits.  Samples SS-1582, SS-1583, SS-1584, and SS-1585 were spiked with 10,000 ug/L, 1,000 ug/L, 50 ug/L and 1 ug/L of each metal, respectively.  The accuracy criteria were met for all of the spiked sample concentrations.

# IX.	ICP Serial Dilution

Thirty serial dilutions were performed on samples from this SDG.  The following samples had recoveries which exceeded the 10 percent acceptance criteria, and with sample results for the metal that were more than 50 times the MDL.  In all cases, the sample result for that metal in that sample were qualified as estimated “J.”

SS-1365	Cr

SS-1395	As and Cr

SS-1416	Cr

SS-1436	As

SS-1456	As, Cr, and Cu

SS-1466	Cr

SS-1566	Cr

SS-1607	As

For samples where one or more metal exceeded the 10 percent criteria, but the result was less than 50 times the PQL, no action was taken.

# X.	ICP-MS Internal Standards

All of the Internal Standard recoveries were within control limits.

# XI.	Field Duplicates

The following primary and blind field duplicate samples were collected as part of this SDG and were collected at a frequency of almost 10 percent:

Primary Sample		Blind Field Duplicate

SS-1587	SS-1324

SS-1588	SS-1333

SS-1589	SS-1344

SS-1590	SS-1348

Primary Sample		Blind Field Duplicate

SS-1591	SS-1352
	SS-1592	SS-1371

SS-1593	SS-1376

SS-1594	SS-1383

SS-1595	SS-1386

SS-1596	SS-1407

SS-1597	SS-1415

SS-1598	SS-1479

SS-1600	SS-1487

SS-1601	SS-1495

SS-1602	SS-1502

SS-1603	SS-1512

SS-1604	SS-1514

SS-1605	SS-1519

SS-1606	SS-1529

SS-1607	SS-1531

SS-1608	SS-1534

SS-1609	SS-1536

SS-1610	SS-1540

SS-1611	SS-1566

SS-1612	SS-1568

SS-1613	SS-1576

SS-1614	SS-1578

All RPD were well below the QC limit of 50% - a reasonable control limit for wipe samples.

# XIII.	Overall Assessment

The reported data meets the acceptance criteria established in the NFG and associated methods and plans. Data quality is sufficient for its intended use.