WHOLE RECORD REPORT( + ADDENDUM)

| MSFC Record # A11207 | In-Flight Anomaly Number -- | Contractor Report Number E-111-2 | JSC# -- | KSC# -- |
| --- | --- | --- | --- | --- |
| Problem Title ULLAGE PRESSURE TRANSDUCER HAD HIGH CONTACT RESISTANCE (SIGNAL DROPOUT) DURING ATP |
| EICN# -- | ELEMENT ET | Contractor MMMSS | FSCM# -- | FCRIT 1R |
| HCRIT -- | Sys_Lvl N | Misc Codes A B C D E F G H I J K L M N O |
| HARDWARE EIM | NOMENCLATURE PRESSURE TRN | PART# NOTED | SER/LOT# NOTED | MANUFACTURER NOTED |
| HARDWARE LRU | NOMENCLATURE N/A | PART# N/A | SER/LOT# N/A | MANUFACTURER N/A |
| HARDWARE NCA | NOMENCLATURE L02 PRESSURE TRN | PART# PD7400098-079 | SER/LOT# 1578 | MANUFACTURER GULTON |
| Test/Operation A - ATP | Prevailing Condtion F - FUNCTIONAL | F / U F | Fail Mode EG - SIG HI OR LO | Cause U - UNKNOWN |
| System ELECTRICAL | Defect -- | Material E - EL C/W | Work Contact J. ADAMS | Fail Date 12/10/1987 |
| Received at MSFC 01/26/1988 | Date Isolated -- | FMEA Reference 3.2.1.2.1 | IFA: Mission Phase -- | Mission Elapsed Time -- |
| Location GULTON | Symptom EE - RANDOM | Time Cycle N/A |
| Effectivity Text LWTS 16, 20, 21, 22, 24 AND SUBS |
| Vehicle Effectivity Codes |
| Vehicle 1 -- | Vehicle 2 -- | Vehicle 3 -- | Vehicle 4 -- | Vehicle 5 -- |
| Mission Effectivity Codes |
| Mssn 1 -- | Mssn 2 -- | Mssn 3 -- | Mssn 4 -- | Mssn 5 -- |
| Estimated Completion Dates |
| MSFC Approved Defer Until Date -- | Contractor Req Defer Until Date -- | LVL 3 Close -- | Remark / Action -- |
| Investigation / Resolution Summary |
| Last MSFC Update 02/13/1995 | CN RSLV SBMT 05/13/1988 | Defer Date -- | Add Date 01/29/1988 | R/C Codes 2 - MFG -- -- |
| Assignee |
| Design P. MULLER | Chief Engineer J. NICHOLS | S & MA R. JACKSON | Project M. PESSIN | Project MGR P. BRIDWELL |
| Approval |
| Design A. JACKMAN | Chief Engineer J. NICHOLS | S & MA R. JACKSON | Project M. PESSIN | Project MGR P. BRIDWELL |
| PAC Assignee J.EL-IBRAHIM | PAC Review Complete JE | MSFC Closure Date 06/03/1988 | Status C - CLOSED | F/A Completion -- |
| Problem Type -- | SEV -- | Program Name -- | REVL -- | OPRINC -- |
| FUNC MOD -- | Software Effectivity -- -- -- -- -- | Software Fail CD -- | SUBTYPE -- | Software Closure CD -- |
| RES PERSON L2 -- | Approval Signature L3 -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Contractor Status Summary |
| Reliability/Quality Assurance Concerns, Recommendations: |
| Problem Description TRANSDUCER HAD HIGH CONTACT RESISTANCE (SIGNAL DROPOUT) AT APPROXIMATELY 9.7 PSID, ON BOTH PRESSURE RISE AND FALL, DURING THE HIGH ACCURACY CALIBRATION TEST. NO SIGNAL DROPOUTS ARE ALLOWED REF. MARS T-34368, CAPS E-082, E-091, E-101, E-106 AND E-110 |
| Contractor Investigation/Resolution CAUSE: THE ROOT CAUSE COULD NOT BE DETERMINED WITH CERTAINTY. A POSSIBLE CAUSE, BASED ON PREVIOUS EXPERIENCE, IS FOREIGN PARTICLE CONTAMINATION OF THE INTERNAL POTENTIOMETER R/C: THE VENDOR HAS CHANGED THE PRODUCTION FLOW TO REDUCE THE LIKELIHOOD OF CONTAMINATION OF THE MECHANISM GENERAL: THE TRANSDUCERS FAILED TO CONTACT RESISTANCE AND SIGNAL DROPOUT CRITERIA DURING VENDOR ACCEPTANCE TESTING (ATP). OCCASIONAL ATP FAILURES OF THESE TYPES ARE EXPECTED AS PART OF NORMAL PRODUCTION FALLOUT. INVESTIGATION OF PREVIOUS SIMILAR FAILURES HAS FOUND THAT MICROSCOPIC PARTICULATE CONTAMINATION OF THE INTERNAL RESISTIVE ELEMENT AND WIPER IS THE MOST PROBABLE CAUSE. REFERENCE: CAPS E-082 THE FAILURE WAS DETECTED DURING VENDOR ACCEPTANCE TEST WHICH IS DESIGNED TO SCREEN DEFECTIVE HARDWARE 5/3/88 - THIS REPORT HAS BEEN DEFERRED FOR STS-26, PER NSTS 07700, VOLUME XI AND NSTS 08126 PARAGRAPH 3.4.1, ITEM C WHICH STATES "PROBLEM CONDITION DOES NOT EXIST IN THE FLIGHT HARDWARE AND IS CLEARLY SCREENED BY ACCEPTANCE TEST, PREFLIGHT CHECKOUT, OR SPECIAL TEST." THE DEFERRAL STATUS HAS BEEN APPROVED BY THE ET PROJECT MANAGER, MR. G.P. BRIDWELL ____G.P. BRIDWELL (SIGNED) 5/13/88____ 5/13/88 TASK I FAILURE INVESTIGATION FAILURE ANALYSES WERE PERFORMED ON THE TRANSDUCERS AT MAF. THE RESULTS ARE LISTED BELOW THE TRANSDUCER HAD FAILED THE VENDOR ACCEPTANCE TEST WITH A LOSS OF SIGNAL, DROPOUT, DURING THE HIGH RESOLUTION CALIBRATION. THE DROPOUT OCCURRED AT APPROXIMATELY 9.7 PSID THE TRANSDUCER INTERNAL MECHANISM WAS REMOVED FROM ITS CASE BY CUTTING AWAY THE CASE WELDS. OPTICAL MICROSCOPE EXAMINATION OF THE INTERNAL MECHANISM AT UP TO 70 DIAMETERS MAGNIFICATION COULD FIND NO CONTAMINATION OR OTHER DISCREPANCIES. THE POTENTIOMETER WIPER TENSION MET THE MANUFACTURERS SPECIFICATIONS. SCANNING ELECTRON MICROSCOPE EXAMINATION OF THE WIPER AND THE RESISTIVE ELEMENT COULD FIND NO CONTAMINATION OR OTHER DISCREPANCIES. THE POTENTIOMETER WIPER TENSION MET THE MANUFACTURERS SPECIFICATIONS. SCANNING ELECTRON MICROSCOPE EXAMINATION OF THE WIPER AND THE RESISTIVE ELEMENT COULD FIND NO CONTAMINATION OR OTHER CONDITIONS WHICH COULD HAVE CAUSED THE FAILURE TASK CLOSED TASK II CORRECTIVE ACTION NO CORRECTIVE ACTION IS REQUIRED. THE VENDOR HAS CHANGED THE PRODUCTION FLOW SINCE THE UNITS COVERED BY THIS CAPS WERE BUILT SO AS TO REDUCE THE LIKELIHOOD OF CONTAMINATION OF THE MECHANISM AT THE TIME THE TRANSDUCERS COVERED BY THIS CAPS WERE ASSEMBLED, THE FOLLOWING CONDITIONS EXISTED: 1. THE OPERATING MECHANISM OF EACH UNIT WAS INSPECTED AT 30X MAGNIFICATION USING AN INCANDESCENT LIGHT SOURCE 2. THE UNITS WERE TRANSPORTED FROM THE INSPECTION AREA TO THE ASSEMBLY AREA, WHERE THE CASES WERE LOOSE INSTALLED, UNDER PLASTIC "MEDICINE BOTTLE" COVERS 3. THE UNITS WERE THEN TRANSPORTED TO THE WELD AREA TO ATTACH THE MECHANISM TO THE CASE THE CURRENT ASSEMBLY STEPS ARE AS FOLLOWS: 1. THE OPERATING MECHANISM AND THE CASE OF EACH UNIT ARE INSPECTED AT 30X MAGNIFICATION ON A LAMINAR FLOW BENCH. THE LIGHT SOURCE IS A FLUORESCENT TYPE WHICH DOES NOT "WASH OUT" DETAILS OF THE COMPONENTS 2. THE CASES ARE LOOSE INSTALLED WHILE STILL ON THE LAMINAR FLOW BENCH 3. THE LOOSE ASSEMBLED UNITS ARE TRANSPORTED TO THE WELD AREA IN A COVERED METAL CONTAINER THESE CHANGES REDUCE THE CHANCE OF CONTAMINATION AT THIS STAGE OF ASSEMBLY WHEN THE OPERATING MECHANISM IS MOST EXPOSED. HOWEVER, THE MANY STAGES OF TESTING, ASSEMBLY OF THE HEADER ONTO THE CASE, AND THE CASE CLOSURE WELDS COULD STILL ALLOW CONTAMINATION TO ENTER THE TRANSDUCER. REFERENCE MEMORANDUM FROM M. FISHER TO C. COYAN, "FOREIGN MATTER IN GULTON PRESSURE TRANSDUCERS." TASK CLOSED TASK III CLEARANCE OF EFFECTIVITIES ALL ETS ARE CLEARED. THE FAILURES WERE DETECTED DURING VENDOR ACCEPTANCE TESTING TASK CLOSED TASK IV CAPS CLOSURE SUMMARY BOTH TRANSDUCERS FAILED DURING VENDOR ACCEPTANCE TESTS. THE FAILURES WERE OF TYPES CONSIDERED NORMAL PRODUCTION FALLOUT. NO CORRECTIVE ACTIONS ARE REQUIRED TASK CLOSED THE CLOSURE IS SUBMITTED FOR MSFC REVIEW AND APPROVAL |
| MSFC Response/Concurrence |

ASSESSMENT ADDENDUM REPORT

| MSFC Report# A11207 | IFA# -- | Contractor RPT# E-111-2 | JSC# -- | KSC# -- | EICN# -- |
| --- | --- | --- | --- | --- | --- |
| Asmnt Part# PD7400098-079 | Asmnt Part Name LOX PRES TRANSDUCER | Asmnt Serial/Lot# 1578 |
| HCRIT CD -- | FCRIT CD 1R | CAUSE CD MAE - MFG-ASY-ENVR | FAIL MODE EG - SIG HI OR LO |
| Asmnt FMEA 3.2.1.1 | Asmnt FM 1 | FMEA CSE E | FMEA SCSE 1 |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Correlated Part# -- | Correlated Part# -- | Correlated Part# -- |
| Associated LRU# -- | Associated LRU# -- | Associated LRU# -- |
| MAJOR DESIGN CHANGES |
| APRV DATE -- | DESCRIPTION OF CHANGES -- |
| ASSESSMENT TEXT |