Fermi National Accelerator Laboratory

Technical Division

SRF Department

P.O. Box 500

Batavia, IL  60510

Fermilab

3.9-GHz Helium Vessel, Low Temperature Maximum Allowable Working Pressure

EDMS # D00000

| Rev | Date | Description | Originated by | Approved by |
| --- | --- | --- | --- | --- |
| None | 13 March 2009 | Final version for operational readiness review | Tom Peterson |  |

3.9-GHz Helium Vessel, Low Temperature Maximum Allowable Working Pressure

Tom Peterson

First draft 2 January 2009, revised 13 March 2009

Introduction.

The pressure vessel engineering notes for the 3.9 GHz RF cavity helium vessels for the tests in Fermilab’s Horizontal Test Cryostat (HTC) established a maximum allowable working pressure (MAWP) in accordance with Fermilab’s ES&H Manual chapter 5031 (following ASME code with some exceptions) of 2.0 bar both warm and cold [ref. 1].  The helium vessels installed in the 3rd harmonic cryomodule are designated as cavity #5 (IND-102), cavity #3 (IND-112), cavity #7 (IND-114), and cavity #8 (IND-115).  Maximum pressure in HTC, even with a cold loss-of-vacuum accident, would not exceed 2.0 bar, so a higher MAWP for cold operations was not necessary.  (“Cold” means cryogenic temperatures of 80 K or less.)

However, in FLASH at DESY, the 3rd harmonic cryomodule will be connected in a system for which low temperature pressures could possibly exceed 2.0 bar, and for which the 1.3 GHz RF cavity helium vessels are rated up to 4.0 bar cold (and 2.0 bar warm).  These 3rd harmonic helium vessels must be rated at a pressure consistent with the rest of the system, a low temperature MAWP of 4.0 bar.

Scaling room temperature MAWP to low temperatures.

The ASME Pressure Vessel Code (referred to as the “code”), Section VIII, Div 1, UG-98(c) says, “Maximum allowable working pressure may be determined for more than one designated operating temperature, using for each temperature the applicable allowable stress value.” [ref. 2]  Thus, one may have a higher cold MAWP with a higher cold allowable stress.

Stress evaluation.

The pressure vessel engineering notes for the 3rd harmonic cavity helium vessels included, as the third load case, “Primary + Secondary stresses evaluation at cryogenic temperature (internal pressure, own weight of the assembly, thermal contraction)”.  (See “3rd load case”, pp. 50 -61, in “Pressure Vessel Engineering Note for the 3.9-GHz Helium Vessel, Cavity #5” [ref. 1].  The stresses were evaluated for 2.0 bar pressure, low temperature material properties, and compared to room temperature allowable.

Multiplying the predicted stress for 2.0 bar by 2 overestimates the stress with 4.0 bar pressure, since it is equivalent to multiplying pressure, gravity, and thermal contraction by 2.  Thus, in the concluding summary table below, 4.0 bar stresses are conservatively estimated as 2 x the 2.0 bar stresses.

Niobium low temperature allowable stress.

The document, “Engineering note: Summary of Niobium mechanical properties” [ref. 3], contains a thorough review of the reported properties of high RRR, heat-treated niobium such as in our 3.9 GHz RF cavities.  The conclusion is that a room-temperature allowable stress for niobium, determined by ASME formulas from very conservatively low ultimate (115 MPa) and yield (38 MPa) strengths, is 25 MPa [ref. 3, pg. 13].  The additional 0.8 derating factor in accordance with Fermilab rules for an “in-house” pressure vessel results in a 20 MPa allowable stress at room temperature, which is the value used in our pressure vessel engineering notes.

The summary of niobium mechanical properties also finds low temperature yield stresses of 320 MPa or greater and ultimate strengths of 600 MPa or greater.  These result in a cold allowable stress based on tensile strength of 600/3.5 x 0.8 = 137 MPa [ref. 3, pg. 13], over six times the 20 MPa room-temperature value.  Since doubling pressure from 2.0 bar to 4.0 bar would no more than double the stresses, the increase of low temperature allowable stress to 137 MPa is more than enough of an increase in strength to allow a cold MAWP of 4.0 bar.

Titanium.

The allowable room temperature stress for grade 2 titanium in our pressure vessel analysis [ref. 1] is, including the Fermilab factor of 0.8, 0.8 x 83 MPa = 66.4 MPa, where the 83 MPa comes from the ASME code allowable stress tables, Tab 1B, Section II, Part D [ref. 2].

The only location where doubling of calculated stress would result in a stress exceeding the room-temperature allowable is in the titanium bellows.

Low temperature ultimate strength for titanium increases to 872 MPa at 77 K [ref. 3, Table 4, pg. 10, for grade 2 titanium, welded] resulting in a primary allowable stress of 872/3.5 x 0.85 (welded pipe) x 0.8 (taking Fermilab 0.8 factor) = 169 MPa, more than double the room-temperature allowable.

Niobium-Titanium.

The room-temperature allowable stress for niobium-titanium alloy (titanium- 45% niobium), determined by ASME formulas from (Wah-Chang data) ultimate stress (546 MPa) and yield strength (480 MPa) is 132.6 MPa.  The additional 0.8 derating factor in accordance with Fermilab rules for an “in-house” vessel results in an allowable stress of 106 MPa.

Room-temperature allowable stress is sufficient to allow doubling of all calculated stresses in the NbTi and still not exceed allowable.  So it is not necessary to invoke a higher low temperature allowable for NbTi.

Cavity collapse.

The non-linear cavity buckling analysis in the pressure vessel engineering note [ref 1, pg. 66] finds a safety factor of 10 using room-temperature niobium properties.  Deformation which may lead to buckling is seen at 20 bar external pressure.  Thus, even using room-temperature properties, we have a safety factor of 5 relative to 4.0 bar on the cavity.

Required test pressure for a 4 bar cold MAWP

From ASME Section VIII, Div 2, Part 8, 8.3.1, the required room-temperature test pressure for qualifying a vessel for a temperature other than room-temperature is Pt = 1.15 x MAWP x (St/S) where Pt is test pressure, S is allowable stress at the operating temperature, and St is allowable stress at the test temperature.  As the tabulated values below show, our 2.3 bar (33.5 psid) pressure test exceeds the required test pressures for the 4 bar cold MAWP.

Conclusion.

|  | Peak cold stress location | Peak cold stress at 2.0 bar x 2 (MPa) | Cold allowable stress (MPa) | Test pressure for 4 bar cold MAWP (bar) |
| --- | --- | --- | --- | --- |
| Niobium | Cavity end cells region | 51.7 x 2 = 103.3 (Cavity 5 note Table A18) | 3xS = 411 | 1.15 x 4.0 x (20/137) = 0.67 |
| Titanium Gr. 2 | Bellows central convolution | 115.6 x 2 = 231.2 (Cavity 5 note Table A17) | 3xS = 508 | 1.15 x 4.0 x (66.4/169) = 1.8 |
| NbTi | NbTi joint, Ti shell weld region | 40.9 x 2 = 81.8 (Cavity 5 note Table A22) | 3xS = 318 (room temperature allowable) |  |

Summary table – worst-case low temperature stresses scaled up to 4 bar internal pressure, compared to cold allowable stress.  Cold stress from the engineering note is multiplied by 2 for a most conservative estimate of stress with 4.0 bar internal pressure.  Note that total allowable stress is 3 x the primary allowable since secondary stresses are also included.

Four bar internal pressure results in low temperature stresses which are less than the low temperature allowable stress for all three helium vessel materials:  niobium, titanium and NbTi.  Thus, the cold MAWP for these helium vessels may be set at 4.0 bar.

References

- Pressure Vessel Engineering Note for cavity #5 (IND-102).  Pressure vessel engineering notes for cavity #3 (IND-112), cavity #7 (IND-114), and cavity #8 (IND-115) are “similar vessels” to IND-102 per FESHM 5031 and have short engineering notes highlighting the differences from IND-102.  These engineering notes have been placed on a public website at http://ilc-dms.fnal.gov/Workgroups/CryomoduleDocumentation/3rd-Harmonic-Cryomodule-for-DESY/  and are also archived with Fermilab’s Technical Division safety office and Fermilab’s ES&H section.

- ASME Boiler and Pressure Vessel Code, 2007 edition

- “Engineering note: Summary of Niobium mechanical properties,” by G. Wu, H. Edwards, T. Peterson, dated 12 December 2008.  Part of the 3rd harmonic cryomodule documentation package located online at http://ilc-dms.fnal.gov/Workgroups/CryomoduleDocumentation/3rd-Harmonic-Cryomodule-for-DESY/ .