(

Test Results of HD1b, an Upgraded 16 Tesla Nb3Sn Dipole Magnet

A. F. Lietzke, S.E. Bartlett, P. Bish, S. Caspi, D. Dietderich, P. Ferracin, S.A. Gourlay, R.R. Hafalia, C.R. Hannaford, H. Higley, W. Lau, N. Liggins, S. Mattafirri, M. Nyman, G. Sabbi, R. Scanlan, J. Swanson

Abstract— The Superconducting Magnet Group at Lawrence Berkeley National Laboratory has been developing high-field, brittle-superconductor, accelerator magnet technology.  A brittle conductor’s support system can significantly impact conductor performance (as well as magnet training).  A recent H-dipole coil test (HD1) achieved a peak bore-field of 16 Tesla, using two, flat-racetrack, double-layer Nb3Sn coils.  However, its 4.5 K training was slow, with an erratic plateau at ~92% of its un-degraded “short-sample” expectation (~16.6 T).  Most of the quench-origins correlated well with regions where low conductor pre-stress had been expected (3-D FEM predictions, 300 K coil size variations).  The coils were re-assembled, with minor coil-support changes and re-tested as “Hd1b”, with a 185 MPa average pre-stress (30 MPa higher than HD1, with a 15-20 MPa pole-turn margin expected at 17 T).  Training started higher (15.1 T), and quickly reached an only slightly higher but stable plateau at 16 T.  A second thermal cycle started at 15.3 T, but, after peaking above 15.7 T, degraded to a 15.6 T plateau.  The temperature dependence of the degraded plateau was explored down to a bath temperature of 3.0 K. Implications for future high-field magnet development are discussed.

Index Terms— Superconducting magnets, Dipole, High-Field, Nb3Sn, Test Results.

# INTRODUCTION

L

awrence Berkeley National Laboratory (LBNL) is continuing a vigorous development program for providing cost-effective, high-field magnet options for “next generation” particle accelerators and storage rings. These magnets utilize state-of-the-art Nb3Sn superconducting cable, and require strong, rigid, predictable mechanical support systems, able to protect the brittle conductor from large Lorentz loads.  While all successful high-field Nb3Sn efforts have thus far utilized wind-and-react technology [1-5], react-and-wind technology is being pursued elsewhere [6].  Previous efforts explored the dual-bore, common-coil geometry, using flat racetrack coil modules [3-5], and achieved 14.5 T bore field.  The most recent effort achieved a bore field of 16 T in a single-bore H-dipole using two double-layer racetrack coils [MT-18]. In all cases, extreme care was taken to avoid conductor damage from excessive pre-stress, tolerating the slow training, to determine how well the conductor and structure responded to large Lorentz loads.  While many flux-jumps were recorded, only conductor movements appeared to initiate quenching, with all models exceeding 90% of “undegraded short-sample”.

HD1’s quench initiations below 16 T were clearly initiated by conductor movements, and its quench origins were strongly correlated with regions where coil size variations [7] and 3D FEM analysis [8] had predicted conductor movement to be likely.  This strong correlation, supported by visual inspection of the disassembled coils, prompted efforts to correct known pre-load deficiencies and retest the coils, to determine if we would make progress, without causing any new problems.  To emphasize the re-use of the same coils and magnetic configuration, this test was called HD1b, and its subsequent thermal cycle test was called HD1b-2.  Their primary results are reported herein.

# Magnet features and test set-up

## Coil inspection and changes

HD1’s coils were disassembled, measured and visually inspected. These 2-layer coils had been wound with 36-strand Rutherford cable, 35 turns (inner layer) and 34 turns (outer layer), using coils 0.8 mm strands of Oxford’s best “restacked-rod processed” Nb3Sn conductor (Jc > 3000 A/mm2 @ 12 T) [9].  Additional details can be found elsewhere [10-13].  A crescent shaped region of crazed epoxy was easily visible at the outer tip of each hard end-spacers (turn #6 away from the pole).  These regions had the approximate shape predicted by our 3D FEM analysis [7,8].  Excepting the high pressure surface impressions left by the G-10 coil-face shims, the coils appeared as virginal as before HD1’s initial assembly: clear epoxy, no visible cracks.  Coil size measurements indicated that coil-A, where HD1’s “plateau” quenching appeared to have been localized to a region that subsequently exhibited a 300K pre-stress deficit, resulting from and under-sized coil support element.  These variations were corrected by judicious application of Araldite-1580 glass-loaded “green-putty”, which was subsequently machined to match the larger coil (coil-B).

## Magnet assembly and Support Structure Changes

The coils were reassembled in the same HD1 geometry (Fig.1) [ray’s and Al’s MT18 papers], where an aluminum shell pre-stressed the coils transversely, via the iron yoke, and four axial Z-rods pre-stressed the coil-ends.  As with previous high-field magnets, most of the pre-stress is only applied while cooling the magnet [12].

Fig. 1, HD-1 cross-section: Two horizontal, double-layer pancake coils were separated by a thin bore-plate, and compressed on all sides by compression pads, bridges, and iron yokes.  Compression was maintained by a 45 mm thick, 740 mm OD tensioned aluminum shell.  G-10 strips electrically insulated the coil surfaces, while simultaneously cushioning irregularities in the high-pressure interfaces.  Iron interference keys permitted removal of internal pre-stressing bladders.  Most of the final pre-stress resulted from cool-down.

Four assembly differences are noteworthy:

1) The shell tension was increased in a manner to produce a ~20% larger cold pressure along the long edge of each coil.  The calculated average coil-edge pre-stress was increased from 155 MPa to 185 MPa, increasing the pole-turn margin to 15-20 MPa at 17 T.  The intention was to stop the pole-turn quenching observed in HD1, while providing an opportunity to observe pole-turn conductor degradation that might result.

2) All coil end-shoes were machined in a manner to more effectively reduce the calculated, Lorentz-induced, coil-end conductor displacements [8].

3) The end-load, from the tension of the four Z-rods, was increased 15% above that used in HD1.

4) The 300K Z-load was applied before the shell tension was raised (in the hopes of reducing the coils’ end-fluff before it was trapped by friction from the transverse loading.

Interface friction in the model was adjusted to obtain agreement with HD1 measurements. As with recent LBNL high-field magnets [4-5, 12-13] the pre-stresses were expected to increase significantly during cool-down (Fig. 2).

## Diagnostic and Test Set-up

The coil protection and diagnostics unchanged from HD1 [13].  Each coil had eleven voltage taps per layer, dividing each layer into 10 segments, with primary attention to the inter-layer ramp, pole-turn, end-spacer turn, outer turn, and lead-in Nb3Sn/NbTi splices.  The average coil pre-stress was inferred from temperature-compensated strain measurements of the aluminum shell, and two of the four axial rods.  A cold Hall-probe measured the peak bore field, with the calibration as used for HD1.

# Test results

## Training Preparations

Cool-down (Fig. 2) supplied most of the pre-stressing tension, as expected, in the shell (30 to 110 MPa) and the Z-rods (125 to 230 MPa).  HD1’s test preparations were repeated [Al’s MT-18 paper].

Fig. 2. The expected absolute shell and axial Z-rod tensions for the following conditions: relaxed onto the keys “293K”, b) cold “4.5 K”, and d) various fractions of short-sample Lorentz excitation.

## HD1b training (1st thermal cycle)

In order to differentiate between hypothetical causes for HD1’s strain-gauge ratcheting, the largest fraction of which was observed after the first quench, the first ramp to quench was changed from our standard stair-step ramp to an ascending series of triangle ramps. This allowed measurements of the ratcheting as a function of peak Lorentz stress (10%, 30%, 50%, 70%, 90% of Iss2), while affording observation of stick-slip training retention and flux-jump repeatability. Most of the fast flux changes have been labeled “flux-jumps”, because they were slow (5-10 ms duration), showed no deceleration (no negative second time derivative of flux), no ringing, and recurred repeatedly every ramp, up and down.

HD1b’s first training quench (10.2 kA, 15 T) occurred while attempting the “90%” hysteresis ramp.  Training was relatively rapid, reaching a ~16 T “plateau” after the 8th quench (Fig. 3). This plateau, while very similar to HD1’s peak field, was considerably more stable than HD1’s.  As observed with HD1’s training quenches, multi-turn, simultaneous quench starts were observed in the starting coil layer, with nearly simultaneous quench-starts in the adjacent layer.  Time-of-flight analysis indicated HD1b’s training quench-origins were at the same coil-end locations as HD1: near an end-spacer apex, one on turn-5 (inside the end-spacer), the others on turn-6 (outside the end spacer; one in the lead-end, the others in the return end; two in coil-B’s inner layer-1, the others in coil-A’s outer layer-2.  After the 8th quench, the quench-origin shifted to a dramatically different location, and a modest ramp-rate dependence was observed (Fig. 4). Compared to HD1, HD1b’s training exhibited a higher initial current, fewer quenches to reach 16 T, no pole-turn only quench-origins, and a stable plateau.

Fig. 3 HD1b’s training started at 15.1 T and reached a stable ~16 T plateau on the 9th attempt, in dramatic contrast to HD1’s training.

Fig. 4 HD1b’s ramp-rate dependence: Quench-origins distinguished training quenches from “plateau” quenches.  Above ~6 A/s, quenches started in coil-B.  While starting slowly, like “short-sample” quenches, and exhibiting ramp-rate dependence, the simultaneous, multi-turn initiations of all plateau quenches appeared to be contradict this conclusion.

Below ~6 A/s, all “plateau” quenches started simultaneously (Fig. 5) in the five turns adjacent to coil-A’s winding pole, in both the high-field layer (layer-1), and the low-field layer (layer-2).  Turn-6 and portions the mid-coil multi-turn segments started simultaneously 1.6 ms after quench-initiation. Compensating for the multi-turn segment’s inductance, and applying the respective turn-6 propagation speeds to its neighboring multi-turn segment, at least 8 outer turns and 9 inner turns participated in the second, simultaneous, 1.6 ms quench-start. Assuming one quenching region per turn, 28 turns (41% of coil-A) had quenches propagating steadily within 1.6 ms of quench initiation.

Plateau quenches also produced propagation signals (Fig. 5) that were near overlays of each other.  The axial time-of-flight to the nearest voltage tap (9.0-9.6 ms to tap “A1R2”) varied slightly with “plateau” current, and corresponds to 176 mm, assuming the average propagation speed observed for turns 2-4 (19 m/s) applies to turn-2.  The quench-quench variation translated to +/- 2 mm. The time-of-flight to the second nearest tap (A1R5) yielded a similar estimate for turn-5.  This places the quench-origin(s) within the high-field straight-section of the magnet, near the return end.

The propagation signals also revealed that the quenches in turns 1-6 started propagating considerably (2-4x) faster than was sustained (Fig. 5).  This was true for both layers, with the longest lasting transients speeds in the 3-turn segment A1R5R2, and the largest speed ratio in turn-6.  Assuming the initial speed burst resulted from degraded margin, the largest degraded region corresponds to distance of 81 mm.

Fig. 5 Plateau quench-origins: Overlay of propagation signals for two plateau quenches (Q12 & Q14) illustrate nearly duplicate quench propagation signals, showing nearly simultaneous multi-turn quench-origins in both the higher-field (A1) layer (lower four signals) and the lower-field (A2) layer (upper four signals).  The displayed signals account for 27 of 28 simultaneous quench-front pairs that exhibited steady propagation 1.6 ms after initiation.

For ramp-rates above ~16 A/s, simultaneous multi-turn quench origins were also observed, but with several differences: 1) Quenches started in coil-B.  2) Simultaneous quench-origins were observed only in layer-1, and only in turns 1-4.  3) Simultaneous quenches in layer-2’s turns 5-9 started 3 ms later than those observed for the slower ramp-rate, coil-A quenching. 4) propagation into the ramp segment from layer-2’s lead-end pole-turn after propagating only 2.5 ms (~31 mm), 5) A layer-1, turn-6 quench passed tap B1L6 (lead-end) after propagating 5 ms.

## HD1b-2 training (HD1b’s 2nd thermal cycle)

Training for the second thermal cycle produced a distinctly different quench history (Fig. 6). Training started with a massive conductor movement at 15.4 T, marginally higher than HD1b, and reached a ~15.7 T “plateau” (~10.6 kA, ~0.93% of undegraded average short-sample) on the second attempt (Fig. 6).  Subsequent attempts resulted in slightly lower, stable plateau slightly above 15.6 T.  Quench propagation signals were near replicas of those observed in the earlier HD1b “plateau” quenching, with two differences: 1) The initial quench-speed surges (immediately after quench-initiation) lasted longer for turns 2-4, and 2) Turn-5’s quench passed tap R5 ~1 ms sooner, implying its origin was ~125 mm inward from R5.  Higher ramp-rate (20-60 A/s) quench-origins appeared to be identical to HD1b’s 16A/s quench-origin: coil-B’s lead-end

Fig. 6. Hd1b trainings:  HD1b-2 (2nd thermal cycle) started higher and reached a ~15.8 T plateau” on the second attempt.  This plateau was ~2% lower than HD1b’s, and degraded with subsequent attempts to ~15.7 T.

All but the second quench produced propagation signals that wear near duplicates every quench, indicating tightly clustered quench-origins (as observed during the previous Hd1b “plateau” quenches).  As with HD1b, HD1b-2’s plateau quench signals revealed a near replication, signal for signal.  This implied a common axial origin, which was near HD1b’s “plateau” quench-origins.  As with HD1b’s “plateau” quenches had multiple nearly simultaneous (within 1 ms) quench-starts, at a current that was relatively stable 15.7 Tesla.  The turn-6 end quenching and pole-turn quenching that characterized HD1 training was was completely absent.  No quenches originated in the ramp, in any of HD1’s four training runs (HD1, HD1-2, HD1b, HD1b-2).

## HD1b-2’s sub-cooled training (T < 4.5K)

In an effort to determine if cooling could increase the 4.5 K “plateau”, the magnet was ramped with a low pressure bath (Fig. 7).  While cooling the bath enabled higher fields, it had no impact upon the 4.5 K plateau current.  All quench-initiations, excepting for the highest quench, were replica’s of HD1b’s “plateau” quenches.  The coldest quench also exhibited simultaneous initiations in coil-A, but in the first three turns at the lead-end.  The observed temperature dependence departed from a slightly degraded-Nb3Sn, HD1b load-line dependence below ~4 K, when the leads were no longer immersed.

Fig. 7. HD1b-2 training, cooled LHe bath dependence:  The conductor was expected to be warmer than the bath when the leads were not immersed (T<~4 K).

# Discussion

## Cool-down & ramping observations

Good agreement with calculations continued to raise confidence in the 3-D stress calculations.  The number and size of flux-jumps during magnet ramping, while disappointing, and requiring desensitization the quench-detection threshold, failed to limit magnet performance in any obvious manner.

## Reduction of HD1’s coil-end training quenches

Hd1b’s coil end-support changes were motivated by the agreement of HD1’s training quench-origins with predicted, Lorentz-induced, conductor displacements near the coil end-spacer.  A delayed onset and large reduction in these displacements were anticipated.  While one can never be certain that full magnet disassembly, repeated coil handling, flexing and a similar reassembly will restore the HD1’s original “virginal” condition, the dramatic improvement in training onset and training speed for HD1b was encouraging.  Improvement had been predicted; improvement was observed; and nothing seriously bad happened.

## Elimination of  HD1’s “plateau degrading” quenches

More uniform coil-A dimensions, and the larger, 180 MPa average coil-edge pre-stress were intended to eliminate coil-A’s pole-turn de-training, while affording an opportunity to observe the impact of conductor strain degradation upon magnet performance.  The intended pole-turn quench-origins that had degraded HD1’s plateau, were completely eliminated, and even allowed a slight increase in the ~16 T maximum field.  Unfortunately, HD1b’s field limit appears to have been conductor-limited, suggesting the possibility that Nb3Sn strain-degradation had been encountered (see further discussion below).

## HD1b’s Quench-Origins

The pattern of HD1b’s training quench-origins was an abbreviated copy of HD1: multi-turn simultaneous quenching near a coil end-spacer, primarily coil-A’s return-end, triggered by stick-slip conductor movement.  HD1b’s “plateau” quench-origins were new, and near replicas of each other.  Five features of the associated “plateau” are noteworthy: 1) All quenches started slowly, in the manner observed in previous short-sample experiences [4].  2) The “plateau” current was only marginally higher than HD1’s maximum, despite the elimination of HD1’s limiting pole-turn quench-origins.  3) Within 1.6 ms of quench-initiation ~41% the coil-A was quenching, quite similar to that observed when triggered by massive conductor movement [5].  4) The associated multiple quench-origins were near replicas, every quench.  6) The plateau current exhibited a slight ramp-rate sensitivity. This immediately raised several questions: 1) What is causing the simultaneous, multi-turn quenching, when all other evidence suggests local loss of conductor margin.  2) Does the plateau current have the expected Nb3Sn temperature dependence? 3) How will the training, and the newly established “plateau”, respond to a full thermal cycle?

## HD1b’s 2nd thermal cycle and sub-cooled training

The massive conductor movement observed prior to the first training quench is the largest ever observed at this test facility.  It is believed to indicate some loss of mechanical training memory, which in conjunction with the large improvement upon the second attempt, was re-established very rapidly.  The observed temperature dependence, in conjunction with the slow-down in quench propagation speed (after initiation), supported the hypothesis that the quenches were originating in degraded conductor regions, consistent with the previously observed (HD1b) slow quench-initiation, quench-origin replication, and plateau ramp-rate dependence.

## “Plateau” quench-origin mystery and challenge

These results leave the following questions unanswered: 1)  How does degraded conductor, presumably locally, strain-degraded conductor, generate simultaneous, multi-turn quench-initiation origins?  2) How many of these “locally degraded” regions exist on the way to 17 T?  3) What is the best method of raising the plateau current?

# Conclusion

The block dipole magnet “HD1” reached 16 T before falling back.  It was disassembled.  The coils were examined in considerable detail, modified slightly in the end, and reassembled with higher pre-stress in those regions that had exhhiited HD1 training difficulties.  HD1’s training quench-origins were either greatly reduced or completely gone.  The observed 4.5 K 16 T “plateau” was stable, and exhibited ramp-rate and Nb3Sn temperature dependences.  However, its quench-initiations involved simultaneous multi-turn quench-origins, in which nearly half the coil was quenching within 1.6 ms.  All quenches appeared to propagate away from regions of degraded margin into regions having considerable margin, leaving a mystery as to what mechanism could locally degrade the margin of so many turns in a manner to cause such observations.

Acknowledgements

This was quite an achievement, albeit only one step along a long development path toward a reliable energy-doubler option. Particular mention goes to the conductor manufacturer (OST), and all the administrators, assistants, and technicians, whose contributions made these results possible.

References

A.D. McInturff et al., “Test results for a High Field (13 T) Nb3Sn Dipole”, Proceedings of the 1997 Particle Accelerator Conference, Vancouver, CA, 1997, p.3212.

A. den Ouden et al, “A 10 T Nb3Sn Model Separator Dipole Magnet for the CERN LHC,” Proceedings of MT-15, 1997, pp. 137-140.

A.F. Lietzke, “Test results for RT1…” VLHC Workshop, FNAL, 2000.

A.F. Lietzke et al.,  “Fabrication and Test Results of a High Field, Nb3Sn Superconducting Racetrack Dipole Magnet”, Proceedings of the 2001 Particle Accelerator Conference, Chicago, June 2001.

A.F. Lietzke, et. al., “Test Results for RD3c, a Nb3Sn...Racetrack Dipole Coil” IEEE Trans Appl. Supercond., Houston, 2002, pp. 1292-1296

G. Ambrosio, et. al., “Fabrication & Test of Racetrack Magnet Using Pre-reacted Nb3Sn Cable” IEEE Trans Appl. Supercond., Houston, 2002, pp1284-1287

S. Mattafirri, et al., “Performance Analysis of HD1: A 16 T Nb3Sn Dipole Magnet”, paper 3LG-01 this conference proceedings

P. Ferracin, et al., “Mechanical Analysis of the Nb3Sn dipole magnet HD1”, paper 3LA03, this conference proceedings

J. Parrell et al., “High Field Nb3Sn Conductor Development at Oxford Superconducting Technology”, IEEE Trans. Appl. Supercond. Vol. 13, no. 2, June 2003, pp. 3470-3473.

R.R. Hafalia et al., “A new support structure for high field magnet”, IEEE Trans. Appl. Superconductivity., vol. 12, no. 1, March 2002, pp. 47-50.

S. Caspi et al., “The use of pressurized bladders for stress control of superconducting magnets”, IEEE Trans. Appl. Superconductivity., vol. 11, no. 1, March 2001, pp. 2272-2275.

RH’s MT-18 paper

AFL’s MT-18 paper

# Footnotes

Manuscript was received 05 October, 2004.

*This work was supported under contract DE-AD03-76SF00098 by the Director, Office of Energy Research, Office of High Energy Physics, U.S. Department of Energy.

All authors are or were with Lawrence Berkeley National Laboratory, Berkeley, CA 94720  (A.F. Lietzke can be reached via e-mail: AFLietzke@lbl.gov, phone: 510-486-4572, or fax: 510-486-5310).