##### QADATA Summary – December 2005

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 100 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 05/01/05 through 12/05/05; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and their impact on new data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/FY05_06charts.htm.  Laboratory and district personnel can also make their own custom QADATA retrieval to parallel their objectives by using the QADATA Application online at  http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Nutrients

Orthophosphate, Filtered,  Colorimetry, Labcode 1974, Parmcode 0671H	4

Phosphorus, Filtered,  (Low-level), Colorimetry, Labcode 2331, Parmcode 0666G	4

Phosphorus, Whole-Water Recoverable,  (Low-level), Colorimetry, Labcode 2333, Parmcode 0665G	5

Trace Metals

Barium, Filtered, ICP-MS, Labcode 1786, Parmcode 1005G	6

Barium, Whole-Water Recoverable, ICP-MS, Labcode 2374, Parmcode 1007D	6

Boron, Filtered, ICP-MS, Labcode 2504, Parmcode 1020G	7

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G	8

Mercury, Filtered, CVAF, Labcode 2707, Parmcode 71890C	8

Mercury, Whole-Water Recoverable, CVAF, Labcode 2708, Parmcode 71900D	9

Uranium, Filtered, ICP-MS, Labcode 1797, Parmcode 22703G	9

Zinc, Filtered, ICP-MS, Labcode 1798, Parmcode 1090G	9

Zinc, Whole-Water Recoverable, ICP-MS, Labcode 2390, Parmcode 1092D	10

Majors

Calcium, Filtered, ICP, Labcode 0659, Parmcode 0915D	11

Fluoride, Filtered, ISE, Labcode 0031, Parmcode 0950B	12

Silica, Filtered, Colorimetry, Labcode 0056, Parmcode 0955C	13

Specific Conductance, Whole-Water Recoverable, Electrometry, Labcode 0069, Parmcode 90095A	14

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Phosphorus, filtered, colorimetry, low-level, parmcode 0666G – Slight positive bias (about 5.6%) starting September 2005.

- Phosphorus, whole-water recoverable, colorimetry, low-level, parmcode 0665G – Slight positive bias (about 6.1%) starting mid-July 2005.

- Mercury, filtered, CVAF, parmcode 71890C – Positive bias (about 5.5%) starting mid-August 2005.

- Mercury, whole-water recoverable, CVAF, parmcode 71900D – Positive bias (about 8.4%) starting late August.

- Uranium, filtered, ICP-MS, parmcode 22703G – Negative bias late August 2005 through mid-October (about 8.6%).

- Zinc, filtered and whole-water recoverable, ICP-MS parmcodes 1090G and 1092D respectively – Negative bias (about 6.5% overall) starting July 2005.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Orthophosphate, Filtered Colorimetry labcode 1974 parmcode 0671H | Corrective actions taken by the NWQL appear to have corrected the bias. Results for samples analyzed mid-October 2005 and later show acceptable bias and variability. No corrective action necessary. |  |  |
| Previous comments for orthophosphate, filtered, colorimetry, labcode 1974, parmcode 0671H |
| From QADATA Summary October 2005 | There is a large negative bias in the data starting mid-August 2005. | New standards have been prepared and concentrations verified by two other instruments; these are now in use as of September 19, 2005. There has been no discernable change in the bias using the new standards. | The NWQL investigation of this bias is ongoing. The investigation includes research into preparation of standards, QC samples, instrument cleaning procedures, and possible sample interference. New QC standards have been ordered and will be put online by the end of October. A new instrument cleaning procedure will be implemented this week. The frequency of QC samples has been doubled to attempt to identify possible cases of sample interference. |
| Phosphorus, Filtered Low-level Colorimetry labcode 2331 parmcode 0666G | All results since September 2005 exhibit a positive bias. The average bias is 5.6% above the target value. This bias is not demonstrated for phosphorus determinations by the regular level methods. | New SPEX standards were put on-line on 10/27/05. The system was rebuilt on 10/20/05. A comparison of calibration standards with the alkaline persulfate line (made from an ERA stock solution) indicates that the low-level standards may be low. However, new SRWS samples prepared on 11/08/05 show values that are very near or slightly below the MPV using the existing calibrants. New batches of ammonium persulfate reagent were prepared 3 times in November. The acidic molybdate reagent is supposed to be good for a year but will be remade and tested to see if that reagent has deteriorated. | A new batch of acidic molybdate reagent was made on 12/07/05 to test the response. If necessary, we will prepare calibration standards from the same ERA stock solution used to prepare alkaline persulfate phosphorus standards by the end of December. |
| Phosphorus, Whole-Water Recoverable Low-level Colorimetry labcode 2333 parmcode 0665G | All results since mid-July 2005 show a positive bias. Results since mid-August 2005 exhibit an average positive bias of 6.1% above the target value. This bias is not demonstrated for phosphorus determinations by the regular level methods. | New SPEX calibration standards were approved on 11/15/05 and 11/29/05. The system was rebuilt on 10/20/05. New SRWS standards were put on-line on 11/17/05 and show values very near or slightly below the MPV. A comparison of calibration standards with the alkaline persulfate line (made from an ERA stock solution) indicates that the low-level standards may be low. New batches of ammonium persulfate reagent were prepared 3 times in November. The acidic molybdate reagent is supposed to be good for a year but will be remade and tested to see if that reagent has deteriorated. | A new batch of acidic molybdate reagent was made on 12/07/05 to test the response. If necessary, we will prepare calibration standards from the same ERA stock solution used to prepare alkaline persulfate phosphorus standards by the end of December. |
| Previous comments for phosphorus, whole-water recoverable, low-level, colorimetry, labcode 2333, parmcode 0665G |
| From QADATA Summary October 2005 | All results received since the last QADATA Summary are above the target value. Starting mid-August 2005, results are, on the average, more than 5% above the target value. | The NWQL agrees with the observation. | New stock standards have been ordered so new working standards can be prepared and put online by October 30, 2005. |
| From QADATA Summary August 2005 | Corrective actions taken by the NWQL appear to have corrected the bias. Results received since May 18 show little bias and little variability. No corrective action necessary. |  |  |
| From QADATA Summary June 2005 | All results received since April 2005 are higher than the target value and exhibit an overall positive bias of 5.8% above the target value. A bias of this magnitude is not evident for phosphorus, filtered (lc 2331). | Working standards prepared and put online April 4th appear to have been prepared from a degrading stock. After high bias was noted by analyst, a new stock standard was ordered. New working standards put online on May 18. | Since the preparation of new standards, the bias seems to have improved. We will continue to closely monitor the line. |

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Barium, Filtered ICP-MS labcode 1786 parmcode 1005G | A negative bias is demonstrated as of September 2005. Results since then are on average about 7% below the target value. The most recently received data, however, show improvement. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL QC charts also indicate no current bias. Since recent data show improvement, the analysis will be closely monitored until the next QADATA summary. |
| Barium, Whole-Water Recoverable ICP-MS labcode 2374 parmcode 1007D | The five most recent results suggest that the previous negative bias has been corrected. Analysis should be closely monitored until the next QADATA Summary is distributed. | The digestion of calibration standards made no significant difference in INBD spike recoveries. | Current NWQL QC charts also indicate no bias. The analysis will be closely monitored until the next QADATA summary. |
| Previous comments for barium, whole-water recoverable, ICP-MS, labcode 2374, parmcode 1007D |
| From QADATA Summary October 2005 | The previously noted negative bias has shifted even lower as of July 2005. Results received since the last QADATA Summary are on average 8% below the target value. This bias is not demonstrated for barium analyses by ICP (labcodes 0641 and 2352). | The analyst noted low bias around mid-July, however, NWQL QC charts indicate the bias is improving. Current calibration standards were prepared on August 24th, 2005. The INBD spike QC chart indicates a -6.4% bias since mid-July. Calibration standards were compared to NIST traceable standards on the ICP and values were within 2%. The analyst has put the calibration standards through the INBD process and is determining whether or not digesting the standards can affect the bias. It will take a few weeks to make a conclusive decision. | Analyst is digesting calibration standards to determine if they are causing a negative bias. |
| From QADATA Summary August 2005 | Results for samples analyzed March 2005 through June exhibit a slight negative bias. The most recently received data, however, do not show this bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Boron, Filtered ICP-MS labcode 2504 parmcode 1020G | Results received for samples analyzed as of mid-September 2005 show marked improvement in variability. All results received since the last QADATA Summary are within 1 f-pseudosigma of the target value. No corrective action necessary. |  | Current NWQL QC charts also indicate that variability has improved. |
| Previous comments for boron, filtered, ICP-MS, labcode 2504, parmcode 1020G |
| From QADATA Summary October 2005 | The variability has been increasing over the last six months. Since late June 2005, six out of 37 results (16%) are outside of +/-2 f-pseudosigmas. | The analyst noted an increase in variability on NWQL QC charts in September 2005 (≈9% RSD). The charts include values both above and below the control limits. It was determined that the high values were probably due to boron carry-over. The low values could be due to a bias caused by the calibration standards. New calibration standards were prepared in mid-September. There have not been enough data points on NWQL QC charts to determine if the new calibration standards have helped with the low bias. | The ICP-MS instrument method was modified on October 3rd, 2005, to increase rinse times when the boron concentration is greater than 100 µg/L in environmental samples. We will closely monitor the boron carry-over to evaluate if the increased rinse-time is sufficient as well as continuing to monitor boron for bias. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | Corrective actions taken by the NWQL appear to have improved the bias. Results from samples analyzed mid-September 2005 and later indicate that the magnitude of the previous bias has decreased. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | Current NWQL QC charts also indicate no bias. The analysis will be closely monitored until the next QADATA summary. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary October 2005 | A positive bias is still present and has increased. Results received since the last QADATA Summary are on average 7% above the target value. This bias is not demonstrated for cadmium analyses by ICP (labcode 0673) or for whole-water recoverable cadmium by ICP-MS labcode 2376). | The analyst noted a positive bias on the NWQL QC charts in September 2005 (≈3%). New calibration standards were prepared in mid-September. Current NWQL QC charts indicate the positive bias has been corrected. | We will continue to monitor the cadmium line for bias. |
| From QADATA Summary August 2005 | Results from samples analyzed June 2005 and later indicate that the magnitude of the previous bias has decreased. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| From QADATA Summary June 2005 | Data indicate a positive bias. Fifty-eight of the last 61 values are above the target value. Results as of February 2005 are on average about 6% higher than the target value. | Online QC charts do not indicate a bias. The calibration standards were run by ICP with a percent difference of 2% or less. SRS T145 was run with a result of 9.8ug/L (limits are 8.51 – 10.15 ug/L). A NIST standard was run with a value of 20.18 ug/L (Theoretical value= 20 ug/L). | We have no explanation for the bias observed in the BQS charts. We will work with BQS to identify the cause of the bias. |
| Mercury, Filtered CVAF labcode 2707 parmcode 71890C | Data indicate a positive bias. Results for samples analyzed as of mid-August 2005 are biased on average 5.5% higher than the target value. | Positive bias was noted on NWQL QC and BQS charts in November. New calibrations standards were prepared on November 28th, 2005. Most recent data indicate bias is improved. | We will continue to monitor the mercury line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Mercury, Whole Water Recoverable CVAF labcode 2708 parmcode 71900D | Data exhibit a positive bias. Results for samples analyzed as of late August 2005 are biased on average 8.4% higher than the target value. | Positive bias was noted on NWQL QC and BQS charts in November. New calibrations standards were prepared on November 28th, 2005. Most recent data indicate bias is improved. | We will continue to monitor the mercury line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Uranium, Filtered ICP-MS labcode 1797 parmcode 22703G | Results for samples analyzed late August 2005 through mid-October show a negative bias. The five most recent results, however, show improvement. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | Current NWQL QC charts also indicate no bias. The analysis will be closely monitored until the next QADATA summary. |
| Previous comments for uranium, filtered, ICP-MS, labcode 1797, parmcode 22703G |
| From QADATA Summary October 2005 | A negative bias is demonstrated as of July 2005. Results since then are on average 7.7% below the target value. | The analyst noted a negative bias on the NWQL QC charts in September 2005 (≈-6%). New calibration standards were prepared in mid-September. Current NWQL QC charts indicate the negative bias has been corrected. | We will continue to monitor the uranium line for bias. |
| Zinc, Filtered ICP-MS labcode 1798 parmcode 1090G | Data still exhibit a negative bias. Results from samples analyzed late August 2005 and after are biased on average 9% below the target value. This bias is not evident for zinc determinations by ICP. | The value for zinc in the vendor stock solution has been determined to be 7% higher than indicated on the certificate of analysis. New calibration standards will be prepared using a new stock solution and verified by the end of December, 2005. All data analyzed since the use of this stock have been within control. No corrective action is necessary on past data. | We will continue to monitor the zinc line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Previous comments for zinc, filtered, ICP-MS, labcode 1798, parmcode 1090G |
| From QADATA Summary October 2005 | Data since July 2005 are biased negative by more than 5% below the target value. This bias is not evident for zinc determinations by ICP. | The analyst noted a negative bias on the NWQL QC charts in September 2005 (≈-8%). New calibration standards were prepared in mid-September. Current NWQL QC charts indicate the negative bias has improved, but is still evident. | We will remake the calibration standards and continue to monitor zinc for bias. |
| Zinc, Whole-Water Recoverable ICP-MS labcode 2390 parmcode 1092D | Data still exhibit a negative bias. Results from samples analyzed August 2005 and later are biased on average 8.7% below the target value. This bias is not evident for zinc determinations by ICP. | The value for zinc in the vendor stock solution has been determined to be 7% higher than indicated on the certificate of analysis. New calibration standards will be prepared using a new stock solution and verified by the end of December, 2005. All data analyzed since the use of this stock have been within control. No corrective action is necessary on past data. | We will continue to monitor the zinc line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Previous comments for zinc, whole-water recoverable, ICP-MS, labcode 2390, parmcode 1092D |
| From QADATA Summary October 2005 | Data since July 2005 are biased negative by about 7.5% below the target value. This bias is not evident for zinc determinations by ICP. | The analyst noted a low bias around mid-July. NWQL QC charts indicate the bias is improving. Current calibration standards were prepared on August 24, 2005. The calibration standards were compared to NIST traceable standards on the ICP and values were within 2%. | Standards do not appear to be the cause of the low bias observed on the BQS chart. We will continue to monitor for bias and look for other causes. |
| From QADATA Summary June 2005 | The previously noted negative bias is showing improvement as of mid-April 2005. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| From QADATA Summary April 2005 | Data since December 2004 are biased negative by 6.3%. All results as of then (21 values) are lower than the target value. This bias is not evident for filtered zinc by ICP-MS. | One internal chart shows a -4% bias for December 2004 only and no bias is observed on all charts for the remaining time period. | We have no explanation for the low bias observed in the BQS charts. We will continue to monitor for the low bias and review with BQS in one month. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered ICP labcode 0659 parmcode 00915D | The bias has improved since mid-September 2005. Since then the overall positive bias is less than 5%. The variability and slight bias still present should be closely monitored until the next QADATA Summary is distributed. |  | Current NWQL QC charts indicate no significant bias or variability. The analysis will be closely monitored until the next QADATA summary. |
| Previous comments for calcium, filtered, ICP, labcode 0659, parmcode 0915D |
| From QADATA Summary October 2005 | The data show an upward shift of about 8% starting mid-July 2005. Results since then are on average 5% higher than the target value. | The ICP instruments were moved to a new location at the end of June. The instrument immediately suffered from an instability caused by erratic argon gas supply. This problem reduced the instruments’ productivity by almost 50%. However, all environmental samples that failed QC were rerun. This instability continued until September 13th. The increased scatter and high bias was noted by the analyst during that time, but since the instability of the instrument due to the gas supply had not been resolved until recently, the direct cause of the high bias was uncertain. The gas supply problem has been solved and new calibration standards were prepared at the end of September 2005. | We will continue to monitor calcium for bias and evaluate the success of the new calibration standards. |
| Fluoride, Filtered ISE labcode 0031 parmcode 0950B | The positive bias (overall about 6% above the target value) is still demonstrated for the blind samples submitted by the IBSP. However, the blind samples submitted cover only a small portion of the analytical range for this method and therefore cannot give an accurate quality assessment of the method’s capability over its entire analytical range. NWQL’s on-line QC standards do not exhibit a positive bias consistently across all concentrations. | The NWQL agrees with the assessment. | We will continue to look for causes for the bias and correct as necessary. |
| Previous comments for fluoride, filtered, ISE, labcode 0031, parmcode 0950B |
| From QADATA Summary October 2005 | Data received since the last QADATA Summary indicate a slight increase in the positive bias. | A vendor service call that optimized the system (new electrodes, software, instrument cleaning and maintenance) in May had some effect on bias and improved blank and carryover performance. New standards were put on-line prior to the service call and slightly improved the bias. A new lot of reagent was put on line in August. | New standards were prepared and put online the last week of September; concentrations are being verified by another instrument. We are verifying instrument and software settings. We will continue to monitor the line for improvement in bias. |
| From QADATA Summary August 2005 | Although a positive bias is still demonstrated (overall, about 6% above the target value), all results since May 2005 are within 1 f-pseudosigma of the target value. Analysis should be closely monitored until the next QADATA Summary is distributed. | We are continuing to see if we can improve the line performance. | The line will be monitored for improvement. |
| From QADATA Summary June 2005 | Data still exhibit an overall positive bias of about 6% above the target value. Almost all results for the last six months are higher than the target value. | New standards prepared early May and new electrode installed May 18. Only two data points on BQS charts since mid May. | Continue monitoring line and discuss with BQS during next meeting. |
| From QADATA Summary April 2005 | Data since mid-November 2004 are biased high by 6.2%. Forty-three out of the last 44 results are greater than the target value. | The NWQL agrees with the observation. | We are ordering new stock standards and will prepare and verify new working standards by April 29, 2005. We also ordered a new electrode to see if the current electrode is contributing to the bias. |
| Silica, Filtered Colorimetry labcode 0056 parmcode 00955C | The positive bias can be attributed to blind samples containing reference material M-164. Data from a study undertaken by the NWQL suggest that this method is unable to resolve an undetermined interference in M-164. Blind samples not containing M-164 exhibit little bias and little variability demonstrating that this method is operating within specifications. | The NWQL agrees with the assessment. | None required at this time. |
| Previous comments for silica, filtered, colorimetry, labcode 0056, parmcode 0955C |
| From QADATA Summary October 2005 | An upward shift in the percent recovery is shown starting June 2005. Overall, recoveries have increased by about 5.5% relative to data prior to June 2005. | During May 2005, the Nutrients Unit tested certified reference stocks to use for calibration standards from three different manufacturers. The overall difference among the three was 5%, and the reference stock with the median value was selected for use starting in June 2005. This change caused the small shift in the values of the IBSP samples. It should be noted that the shift in values was smaller for samples that do not contain M164 than for those that contain M164, and that the online third party QC samples that had a small low bias shifted to no bias after the calibration standards were changed. | The NWQL will continue to monitor the Online and IBSP Control Charts. All IBSP mixes that contain M164 have had a high bias for labcode 0056 no matter what the SiO2 concentration of the mix was. Samples that do not contain M164 give values that are close to their MPVs. When samples containing M164 are diluted x10, the interference is much lower. The Nutrients Unit will continue to attempt to determine the cause of the interference. |
| Specific Conductance, Whole-Water Recoverable Electrometry labcode 0069 parmcode 90095A | Corrective actions taken by the NWQL in early August appear to have corrected the bias. Data received since the last QADATA Summary show no bias. No corrective action necessary. |  |  |
| Previous comments for specific conductance, whole-water recoverable, electrometry, labcode 0069, parmcode 90095A |
|  | A downward trend in the percent recovery is shown starting June 10, 2005. Results since then are on average 7% lower than the target value. | A new combination electrode was installed 7/14/05 for pH and conductance. Due to poor performance the combination electrode was replaced with separate electrodes on 7/20/05. Specific conductance QC problems started about 7/26/05. The specific conductance probe was again replaced 8/3/05. The downward trend seen in the BQS charts was observed during the transition to the new probes. Present QC is acceptable and consistent; third party check sample results have been more consistent and close to accepted values. | No more corrective action will be taken at this time. We will continue to closely monitor performance of the specific conductance probe. |