##### QADATA Summary – September 2004

The Inorganic Blind Sample Project (IBSP) reviews 109 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 02/01/04 through 08/26/04; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples that are known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the NWQL which gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The lab’s comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report to help provide a clearer perspective of the overall concern.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY04_05charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Arsenic, Filtered, ICP-MS, Labcode 2503, Parmcode 1000D  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Chromium, Filtered, GFAA, Labcode 1936, Parmcode 1030I . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Strontium, Filtered, ICP-MS, Labcode 2507, Parmcode 1080I . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Uranium, Filtered, ICP-MS, Labcode 1797, Parmcode 22703G. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Sodium, Filtered, ICP, Labcode 0675, Parmcode 0930C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Sulfate, Filtered, IC, Labcode 1572, Parmcode 0945G. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Filtered ICP-MS labcode 2503 parmcode 1000D | The results received since the last QADATA Summary indicate that corrective actions taken by the lab were effective. |  |  |
| Previous comments for arsenic, filtered, ICP-MS, labcode 2503, parmcode 1000D |
| From QADATA Summary July 2004 | A dramatic shift in the data is evident starting mid-May 2004. | Arsenic bias was traceable to the correction factor algorithm and the likely source was in the software. This error caused arsenic values to be biased high. The solution used to check the accuracy of the correction equation did not show the error. | The software was reinstalled by the instrument vendor, thus restoring the correct correction factor algorithm. A hard disk diagnostic check on the instrument computers was performed and no problems were found with disk memory. All data affected by this problem is in the process of being updated and a RapiNote is being drafted. The correction check solution has been changed to include a known amount of arsenic to verify the correction factor is correcting the target arsenic concentration. The SOP will be amended, and training will be documented. We anticipate having the documents and training completed by July 30, 2004. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | The long-running slight positive bias shows improvement. As of June 2004, all results are within 1 f-pseudosigma and data are biased less than 3%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary July 2004 | The long-running slight positive bias (6%) is still evident. This bias is not demonstrated by any other of the five analytical methods analyzing for cadmium. | Internal QC and NIST, both third party checks, indicate a bias of about 2%, which is within the acceptance criteria for calibration standards. Concentration of internal QC and NIST checks has been verified by alternate methods. An internal QC review of the other five methods also shows a 2% or less bias, again within the SOP criteria. Interferences are not a cause for the high bias. | The stock standards are due for the annual replacement at this time. New standards are on order. The process for preparation and review of new standards will follow. |
| From QADATA Summary May 2004 | As of March 2004, results indicate the positive bias is less than 5%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL continues to closely monitor analysis. |
| From QADATA Summary March 2004 | Data from January 2004 through March exhibit a slight positive bias (6%). Almost all results are greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | Due to the on-going monitoring of internal QC, a positive bias is still not evident on NWQL QC charts. To verify this, an additional third-party check standard from a different source than is routinely used by ICP-MS was run during the week of March 8th, 2004, to verify the ICP-MS calibration curve. Calibration and QC are acceptable. | The issue of possible interferences that could affect cadmium is being discussed with the NWQL’s Method Research and Development Program. More research needs to be done. Also, internal QC and IBSP charts will continue to be monitored for bias and variability. |
| From QADATA Summary January 2004 | Data for the last six months (June 2003 through December) demonstrate a slight positive bias (6%) with all 48 values greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of December 2003. We will continue to monitor internal QC and IBSP charts for bias and variability. If the third-party check standards indicate a bias, calibration standards will be remade and verified. |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030I | The chart shows a positive bias (9.3%) beginning June 23, 2004. Since this time, all results (19 values) are higher than the MPV. | NWQL QC charts do not indicate a bias. Eleven of the BQS samples were run on the ICP and all values were higher than the MPV. The QC for the run (CCVs and TPCs) was running slightly lower than the MPV. A NIST traceable standard was run on the GFAA. The values came out slightly lower than the MPV. | NWQL will continue to closely monitor analysis. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | Results received as of mid-June 2004 indicate increased variability. Six out of 27 values (22%) are outside of –2 f-pseudosigma. The data also exhibit a slight negative bias (4.3%). | A software problem was detected in mid-August that caused a slight increase in variability for strontium. Also, at this same time, it was determined that the instrument detector needed replacement. NWQL QC charts indicate a negative bias of 5%. | The instrument vendor has fixed the software problem and a new detector has been installed. Calibration standards will be remade and verified by September 14th. NWQL will continue to monitor strontium for bias and variability. |
| Uranium, Filtered ICP-MS labcode 1797 parmcode 22703G | The chart shows a negative bias (6.2%) beginning July 23, 2004. Since this time, 14 out of 15 results are lower than the MPV. | NWQL QC charts indicate a negative bias of <3%. Analysis of calibration and QC standards against NIST standards show a <3% bias. | New standards are due to be remade and verified by September 14th. NWQL will continue to monitor uranium for bias. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Magnesium, Filtered ICP labcode 0663 parmcode 0925C Sodium, Filtered ICP labcode 0675 parmcode 0930C | The variability has improved to within acceptable levels for both magnesium and sodium (<2% and <3% RSD, respectively) since the last report. Results for sodium indicate no analytically-significant bias (<3%); however, data for magnesium (since May 18, 2004) show a slight negative bias (5.2%) with 43 out of 44 values below the MPV. | NWQL charts indicated a negative bias of 4% for magnesium in mid-August. New standards were made and verified at the end of August. | NWQL will continue to monitor for bias. |
| Previous comments for calcium, magnesium, and sodium by ICP |
| From QADATA Summary July 2004 | Data received since the last QADATA Summary indicate an improvement in the bias for each analyte (all less than 4%). The variability has improved for calcium and is now within acceptable levels; however, magnesium and sodium data still exhibit high variability. | Based on internal QC and instrument optimization procedures the variability mentioned by IBSP is not observed. Contamination, calibration curve design, nebulizer conditions and performance, and sample stability for the longer analytical times have all been examined and proven not to be contributors to the variability concerns. The instrument meets the manufacturers specifications for performance. The method meets or exceeds the requirements for ASTM, Standard Methods, and EPA published methods for performance. | We have nothing new to try at this time. However, we will continue to consider the 4% variability and try to make improvements. |
| From QADATA Summary May 2004 | Data received since the last QADATA Summary indicate a positive bias. In addition, more values have been received that are at or outside of +/-2 f-pseudosigmas. Results from the new instruments (put into service June 2003) still indicate high variability relative to results from the old instruments. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last QADATA Summary (March 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. Tests have been performed that rule out non-instrument related contamination. Tests are being conducted to determine whether instrument contamination and carry-over is present. The positive bias for calcium and magnesium on BQS control charts is also evident on NWQL QC charts, prompting the preparation of new standards. The positive bias for sodium is not evident on NWQL control charts. | NWQL continues to investigate the cause of the outliers on runs with acceptable QC by running tests on the instrument and determining whether contamination and carry-over is present. New standards for calcium, magnesium, and sodium have been prepared and were validated on May 12th, 2004. NWQL will continue to closely monitor the analysis for bias. |
| From QADATA Summary March 2004 | New instrumentation was implemented in June 2003. A comparison of results from the old instruments (TJA) to results from the new instruments (PE) indicates an increase in variability. The OWQRL is not demonstrating variability problems for these constituents. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last data summary (January 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. | As stated in the January 2004 QADATA Summary for cadmium WWR GFAA (labcode 1555), use of an additional QC check sheet that is included in all data packets was implemented to help alleviate incorrect data going in LIMS. This practice was implemented for ICP processes in October 2003, and data shows that it has been useful in eliminating some errors. We will continue to investigate the cause of the outliers on runs with acceptable QC. |
| Sulfate, Filtered IC labcode 1572 parmcode 0945G | The chart indicates a slight negative bias (4.7%) starting early-April 2004. All 29 results since this time (except one outlier) are below the MPV. | We reviewed internal QC to help evaluate the observed bias. Our data indicates a bias of about 3% or less, in only the midsection of the analytical range. We also noticed increased variability during the period of the observed bias. We are currently implementing new instruments on this analytical line, and will eliminate use of the instrument producing the observed bias in the next few months. | We have no explanation of the bias, or why it is only evident in the middle of the analytical range. We will continue to monitor the situation and attempt to find the cause. |