Calibration Results of 60cm Scales

(Jan 2009)

| Scale | Ideal Value |  | Overall | Scale Correction (1/2 overall difference) |
| --- | --- | --- | --- | --- |
| Letter | Number | 0.025400 |  | Difference | scale is ground short |
|  |  | parallel bar |  |  | scale is long |
|  |  | [m] |  | [m] | [m] | [inches] |
| A | 9408 | 0.024890 | Scale is ground short | -0.000510 | -0.000255 | -0.01004 |
| B | 9409 | 0.024920 | Scale is ground short | -0.000480 | -0.000240 | -0.00945 |
| C | 9407 | 0.025070 | Scale is ground short | -0.000330 | -0.000165 | -0.00650 |
| D | 9406 | 0.025385 | Scale is ground short | -0.000015 | -0.000008 | -0.00030 |
| E | 9500 | 0.025385 | Scale is ground short | -0.000015 | -0.000008 | -0.00030 |
| F | 9501 | 0.025235 | Scale is ground short | -0.000165 | -0.000083 | -0.00325 |
| G |  | 0.025285 | Scale is ground short | -0.000115 | -0.000058 | -0.00226 |
| H |  | 0.025345 | Scale is ground short | -0.000055 | -0.000028 | -0.00108 |
| J | 68 | 0.025455 | Scale is long | 0.000055 | 0.000027 | 0.00108 |
| CFR1 |  |  |  |  | 1.064720 |  |
| CFR2 |  |  |  |  | 0.980757 |  |

Matching Scales:

|  | Scale Correction | Difference |
| --- | --- | --- |
|  | [m] | [m] |
| D | -0.000008 |  |
| E | -0.000008 | 0.000000 |
| A | -0.000255 |  |
| B | -0.000240 | 0.000015 |

Measurement of Carbon Fiber Rods in the Garage

Request # 76											02-OCT-2008

Project   # 76											C. Wilson

Completed the measurement of 2 Carbon Fiber Rods (CFR’s) used for precise level runs in the Tevatron. The scale holders on both CFR’s were replaced with new scale holders

and secured using Loctite “Blue”.

New values for the CFR’s

CFR #1    1.064720 Meters   or   41.918104 Inches

CFR #2    0.980757 Meters   or   38.612497 Inches

# Crew Reports

Request #  : 5732                       Date   : 26-Dec-08

Project #  : 1802                       Crew # : 2

Area       : OFFICE

Log Book # :                            Page # :

File Backup To Diskette (Y/N) :         Vax Filename :

Job Completion (Y/N)  : N

Estimate remaining time to finish the request in hour  :   8.0

Description

Collected data for calculating the correction on 60cm scales

A,B,E,F,H, and J. This data was colllected using the DNA03

levels program.

------------------------------------------------------------------------------

Request #  : 5732                       Date   : 31-Dec-08

Project #  : 1802                       Crew # : 7

Area       : OFFICE

Log Book # :                            Page # :

File Backup To Diskette (Y/N) :         Vax Filename : NONE

Job Completion (Y/N)  : N

Estimate remaining time to finish the request in hour  :  20.0

Description

continued testing the NEW levels program, utilizing the program to

also perform 60cm scale references (calibrations for corrections)

------------------------------------------------------------------------------

Request #  : 5732                       Date   : 2-Jan-09

Project #  : 1802                       Crew # : 2

Area       : OFFICE

Log Book # :                            Page # :

File Backup To Diskette (Y/N) :         Vax Filename :

Job Completion (Y/N)  : N

Estimate remaining time to finish the request in hour  :  10.0

Description

More level program testing. Checked 60cm scale C linearity.

------------------------------------------------------------------------------

Request #  :   76                       Date   : 7-Jan-09

Project #  :   76                       Crew # : 3

Area       : CALIBRATION ROOM

Log Book # :                            Page # :

File Backup To Diskette (Y/N) : N       Vax Filename :

Job Completion (Y/N)  : N

Estimate remaining time to finish the request in hour  :   5.0

Description

Checked correction number accuracy on 60 cm scales A,B,C,

E,F,G,H and J. Measured D and K for calibration purposes.

Still need to measure L and M for calibration.

------------------------------------------------------------------------------

Request #  :   76                       Date   : 8-Jan-09

Project #  :   76                       Crew # : 3

Area       : CALIBRATION ROOM

Log Book # :                            Page # :

File Backup To Diskette (Y/N) : N       Vax Filename : 76_08JAN09.DIR

Job Completion (Y/N)  : Y

Note is checked (Y/N) :             Checked by :

Note or file is put inside the basket (Y/N) :

Description

Completed verification checks of 60cm scale correction numbers

for scales A,B,C,E,F,G,H and J. Also, completed re-measurement

of scale D and measurement of new scales K,L, and M for calib.