Line-by-line comments  4-apr-08

16 	The ID does NOT provide charged particle track reconstruction (it would be great if a detector could do that and output tracks rather than hits, but haven’t yet figured out how)

35-50 	Replace “excellent”, “good”, “very good” with numbers where possible, and simply remove where not possible. For example “good 3D vertexing” should become “3D vertexing”. “Excellent transverse impact” should become “order 10um transverse impact”, etc.

43 	and secondary interactions -> , secondary interactions, and degradation of calorimeter energy resolution.

46 	or 10^15 units -> and 10^ units (to be referred to as the lifetime dose)

48	the requirement … confirmed by -> this is based on

53 	smallest -> smallest achievable

55	dose -> ionizing dose

55 	500kGy -> the lifetime dose

67	monitoring and cooling -> monitoring, optical I/O, and cooling

Table 1: 	Number of pixels -> Number of channels                       (not the same thing!)

Table 2: 	ditto

113: 	I would remove sentence “the required radiation tolerance…”

Fig 5 	caption: add “Sensor bias supply not shown”

120 	no new paragraph.

120 	the cells of the sensor array are collected -> the sensor is integrated

123 	serial star -> star            (what is a serial star?)

134 	system uses commercial components -> system uses a combination of customized commercial and fully custom components.

134 	delete “adapted to…pixel detector”

137 	maximum -> absolute maximum              (this is a standard term)

138 	To interface between…implemented -> The optical links are custom made using commercial diode and laser array bare die with custom driver integrated circuits and packaging.

153 	The column readout -> The FE-A column readout

154 	Hits are -> Hits were

160-165 	I have a big problem with such vendor references. I strongly suggest to remove all vendor names from the paper. Otherwise you need to get permission in writing from each vendor, and in any case you surely won’t get it from some so some vendors will be mentioned and not others. I see no possible value in including vendor names and product part numbers in such an overview paper. This should be called a “radiation qualified 0.8um CMOS technology”

167 	left the collaboration with -> pushed the collaboration towards

Table 3: 	Add FE-H before FE-I1. 0.8um radiation qualified CMOS

Figure 6	caption: Floor plan -> Schematic diagram.

Figure 6 	caption: add “Not to scale”.

188 	DC feedback -> feedback capacitor with a current source continuous reset

188 	mA -> uA

190 	bias current of the feedback –> continuous reset

190 	delete “,used to discharge…capacitor”

190 	The feedback system -> The implementation

195 	a pulse width -> a discriminator pulse width

196 	over-Threshold, -> over-Threshold (TOT),

200 	comparator -> discriminator

201 	Comparator -> Discriminator

229 	cells with data -> occupied cells

233 	to transfer a single hit -> at which bytes can be transferred

239 	at a speed -> at a selectable speed

246 	Hit information is -> Once

246 	, and waits there -> , hit information waits there

247 	at the correct time…for the hit -> at a time corresponding to the LE time stamp plus a programmable trigger latency

249 	received triggers -> received one or more L1 triggers

250 	readout sequence in which… -> readout sequence, as described under the next heading.          (move rest of para. to next heading)

255 	Keep first sentence, delete rest of this paragraph (redundant with next heading).

270 	Add new paragraph: During the readout sequence the EOC buffers are scanned…    (paragraph moved from previous heading).

283 	The noise is 160e -> The noise with the sensor load (50um x 400um pixel) is 160e

285 	depends -> depend

290 	600kGy -> the lifetime dose

292 	pixels) -> pixels at -7C)

294 	12 inch -> 8 inch

297 	-> The wafers were probed using custom setups on semi-automatic probe stations. Each chip was fully exercised including measurement of a reference calibration capacitor value using a dedicated circuit only available at probing. This removed process variations from the charge and threshold calibration scale. Test time was approximately 30 hours per wafer and the selection was based on cuts on over 30 analog and digital parameters.  [reference- I think Aldo may have a conference proceeding we can reference- should ask him]. Every chip was probed again after dicing the bumped wafers using custom vacuum chucks to hold up to 60 chips at one time that could be probed like a wafer. The failure rate of chips after dicing was 2% (this would have translated into 30% module rework rate if this step had been skipped). [reference: here we should reference a conf. proceeding or paper by Igor Volobouev, who developed this technique for CDF, in addition to what we may have on it (maybe Bonn)]. Additionally, the indium bumped chips were photographed during probing and 160 images per chip (containing all the bumps at 50X magnification) were archived for future reference.

362	add footnote for LVDS type signals “Low voltage differential but not necessarily conforming to the LVDS standard”

368	2-bit wide serial -> 2-bit parallel

373	add: “the 40MHz beam crossing clock using multi-drop LVDS”

379	command, -> command (11101),

386	the circuit were -> the circuit was

387	Upset -> Upset (SEU).

393	delete “, the total power is 290mW”       (anyone reading this far can multiply 145x2 in their head)

395	-> The wafers were probed at a commercial vendor using ATLAS supplied test vectors. The chip includes special outputs to be able to test the full state machine “phase space” with such vectors.           (Giovanni should rephrase this).

825-830	This is a horrible introduction that says noting at all. Can you simply remove it?

832	resistivity bulk -> resistivity n-type bulk

836	all charges… -> charge carriers generated by ionizing particles passing through the active volume.

836	concept -> design

838	The pixel sensor…n-bulk -> The n-bulk thickness is 250um and the resistivity between 2 and 3 Ohm.cm

838	250um thickness to the checked- I think it’s more like 260. Actual value with errors would be preferable.

838	“400x50um^2” this is impossible- it violates unitarity. The actual size of the implants should be given, with two values for length corresponding to 400 and 600 um pixels

839	implantations -> implants

840	88.9% -> 8/9

841	11.1% -> 1/9

842	implantations -> implants

852	in the bulk … against -> have been introduced into the bulk to increase tolerance against

872	delete “demands precise mask steps”

873	micrometers. -> microns, which makes the manufacturing process demanding.

875	The 250um thick…3Kohm/cm -> The sensors

949	1.86 -> 1.64

949	100um -> ~100um

949	foil -> (“Flex hybrid”)

951	foil -> circuit

956	~190um -> 195 +/-10 um (sigma)

956	by back-side -> by wafer back-side

959	flexible-hybrid kapton foil -> Flex hybrid

960	foil -> Flex hybrid

961	flex-kapton circuitry -> Flex hybrid

963	dispersion on the flex circuit -> variation

964	The flex circuit…sensor bias -> The back side of the Flex hybrid must be pinhole free because it lays directly on top of the high voltage side of the sensor. A multiple solder mask layer was therefore used and all parts were tested to hold of 1000V.

965	The kapton material … (500KGy) -> Since all module components must withstand the lifetime radiation dose, polyimide was used as the Flex hybrid base material with adhesiveless metalization.

968	(NTC) resistor placed -> (NTC) thermistor loaded.

970	A complete … 500 KGy. -> After lifetime dose a module is expected to draw 1.3A at 1.7V analog and 0.9A at 2.1V digital currents, including pigtail and Flex hybrid but not microcable voltage drops,  plus 1mA at 600V sensor bias, for a total power of 4.7W. However voltages may have to be increased to recover performance up to a maximum of 6W.

977	The sensor pixels … (about 11%) which -> 8/9 Of the sensor pixels have dimensions of 50 um x 400um, while 1/9

980	These off-side sensor pixels are connected… 4+4 -> Each of these special pixels is ganged through metal lines on the sensor  with one of 4+4

981	delete “in addition to … underneath”

983	These special pixels … increased sensor capacitance -> Thus there are 5 distinct pixel types with decreased performance due to added input capacitance and ambiguities (see 6.6). These are: long (10.6%), ganged (2.2%),  inter-ganged (1.6%), long ganged (0.3%), and long inter-ganged (0.2%).

992	get -> be

992	for 20cm integrated circuit wafers -> for FE-I3 wafers

995	The integrated circuit wafers … flip-chip process -> The dicing was performed before thinning by making 250um deep cuts in the 700um thick wafers. Separation of the chips occurs towards the end of the back-grinding operation (dicing before grinding). This was done to obtain high quality diced edges.

1002	“which provides self alignment” I have doubts about this since force is used to press the chip and sensor together. I don’t see how force and self alignment are compatible. Should ask Norbert.

1009	The bump pitch -> No under bump metalization is needed. The bump pitch

1017	television -> video

1021	solder and indium -> solder (45%) and indium (55%)

Fig. 25 	not the best. Why not put the SEM x-section of soldered bumped chip instead?

1040	once again remove footnote listing vendor name. Are we going t list all vendors or just selected ones?

1049	zero insertion force -> 30-pin surface mount                  (not ZIF by a long shot)

1051	use thin copper…which are -> have the micro-cable

1056	the glue is deposited -> silicone strip adhesive was used

1058	remove vendor reference 11 (it’s completely wrong anyway)

Fig 28	This is not a picture! It is the same Daniel Dobos CAD model shown in fig 23! I can send you a good picture of a disk module (since fig 23 shows the barrel module a disk module might be appropriate).

1065	A ranking value…of the detector -> This was done by ranking the modules as explained later, rather than using a cut bases selection and treating all selected modules as equal.

1066	In this calculation different weights are assigned to quantities such as the number of missing or shorted bumps, minimum digital voltage for error free operation, sensor bias current, number of reworked wire bonds, etc.

1070	is done correctly -> produced test bonds exceeding a minimum pull force (every module had a number of spare bonds that were pull tested).

1096	A problem in the potting -> A systematic problem in the encapsulation (remove vendor footnote! The problem was not due to this product anyway!)

1097	which was corrected during production ->  . Correcting this problem required the rework of a number of modules, and such a rework carried a ranking penalty.

Fig.29	The thresholds are in fact not uniform for the entire module as can be seen form the fact that the fit is a horrible match to the data. The data look in fact like the sum a few Gaussians with slightly different means. One should probably fit to two Gaussians with ratio of areas fixed to 1:8 (the ratio of normal pixels to other pixels).

1118	The number of dead channels… this function includes: -> Modules were ranked using a single value calculated from the test results. In this calculation different weights are assigned to each variable considered, dead channels from source data carrying unit weight. The information combined in the ranking value includes:

1120+	add bullet: “Number of bad channels, including source scan results, digital tests, and other analog measurements.”

1122	remove “penalties for”

1124	any repair operations performed, including number of reworked wire bonds

1125	function -> value

1125	remove “of the ranking function”

1126	bump -> excess

1126	remove “in the ranking distribution”

1127	potting -> encapsulation

Fig 31	IZM -> solder,  SELEX -> indium

Table 6	yield -> fraction

Table 7	ditto

1203	600V -> 600V sensor bias

Table 8	Do we have to mention ON and CIS? Can we say Vendor A and Vendor B instead?

1280	operating -> sensor

1288	Interpretation of depletion in heavily irradiated silicon is tricky. Can a sensor expert comment look at this and see if there is a different phrasing with reference to sensor paper that is more appropriate?

1290-1549 impeccable

1560	photography -> photograph

1609	additional -> dedicated

1609	keep them at -> actively regulate their temperature at any chosen value up to 30C.

1612	through -> additional

1633	occupancy values need units

1635	ditto