Cruise: EX0645S

Dates:  November 4 - 12, 2006

Chief Scientists:  not applicable

Equipment:  Surface samples collected.

Total number of stations:  VOS Underway Cruise

Chemical measurements:

DIC:  10 locations, 20 samples each 500-ml, 20 duplicate samples

PI:  Dr. Rik Wanninkhof

Collected by:  Kevin Sullivan

Analyzed by:  Esa Peltola

Salinity:  Salinity from the ship’s thermosalinograph on the Practical

Salinity Scale

TALK:

PI:  Dr. Frank Millero

Collected by:  Kevin Sullivan

Analyzed by:  Alex Abrams

Details:

DIC:

Analysis date: December 4, 2006

Coulometers used: AOML1 and AOML2

Blank range: 13.4-28.6 counts/min

CRM # used and assigned value (include both DIC and salinity): Batch 59, c: 2007.1 umol/kg,S: 33.316

CRM value measured:  AOML 1: offset 10.6 umol/kg (2017.7 umol/kg)

AOML-2: offset 10.1 umol/kg (2017.2 umol/kg)

Average run time, minimum run time, maximum run time:  10 min, 8 min, 17 min

Reproducibility: (# samples and average difference): 10 sets of duplicate samples, average difference 1.6 umol/kg

CRM, salinity and HgCl2 correction applied: Salinity correction was applied for the AOML1 samples, the AOML2 samples were calculated using TSG salinity; CRM and HgCl2 volume correction was applied

Location of analyses files: Esa's computer file/directory/name: My Documents/SOMMA/Lab Studies/2006 (12) Explorer /

Remarks-

The volume correction was applied due to added HgCl2 (Measured DIC*1.00037).

The first CRM of each cell was used for a CRM correction.

There was a good agreement between the duplicate samples All 20 samples were sent to Millero's lab for alkalinity analysis