##### QADATA Summary – August 2006

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 100 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 01/01/06 through 07/27/06; however, longer-range charts may be used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the NWQL are removed from the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and how the corrective actions have impacted the new data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/FY06charts.htm.  Laboratory and district personnel can have their own custom QADATA retrieval to parallel their objectives by contacting Ted Struzeski (303.236.1872, struzesk@usgs.gov) or by using the QADATA Application online at  http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Nutrients

Phosphorus, Filtered,  (Low-level), Colorimetry, Labcode 2331, Parmcode 0666G	4

Trace Metals

Chromium, Filtered,  ICP, Labcode 0722, Parmcode 1030E	5

Mercury, Filtered,  CVAF, Labcode 2707, Parmcode 71890C	5

Mercury, Whole-Water Recoverable,  CVAF, Labcode 2708, Parmcode 71900D	5

Nickel, Filtered,  ICP-MS, Labcode 1795, Parmcode 1065G	5

Strontium, Filtered,  ICP-MS, Labcode 2507, Parmcode 1080I	6

Zinc, Whole-Water Recoverable,  ICP-MS, Labcode 2390, Parmcode 1092D	6

Majors

Alkalinity, Filtered, Electrometry, Labcode 2109, Parmcode 29801A	7

Alkalinity, Whole-Water Recoverable, Electrometry, Labcode 0070, Parmcode 90410A	7

Fluoride, Filtered, ISE, Labcode 0031, Parmcode 0950B	8

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Chromium, filtered, ICP, parmcode 1030E – Positive bias (16%) starting May 2006.

- Mercury, filtered, CVAF,  parmcode 71890C – Slight negative starting mid June 2006.

- Mercury, whole-water recoverable, CVAF,  parmcode 71900D – Slight negative bias starting early June 2006.

- Nickel, filtered, ICP-MS, parmcode 1065G – Positive bias starting mid May 2006.

- Zinc, whole-water recoverable, ICP-MS, parmcode 1092D – Slight negative bias starting April 2006.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Phosphorus, Filtered Low-level Colorimetry labcode 2331 parmcode 0666G | Variability has improved and very little bias is present. Most results received since the last QADATA Summary are within +/-1 f-pseudosigma of the target value. No corrective action necessary. |  |  |
| Previous comments for low-level phosphorus, filtered, colorimetry, labcode 2331, parmcode 0666G |
| From QADATA Summary June 2006 | A positive bias starting in 2005 is shown improving in February 2006. However, variability has increased as of late January 2006. Since then, four out of 34 results (12%) are more than 4 f-pseudosigmas lower than the target value. | One of the samples was determined to be a mispour by the analyst. Further investigation is now under way to determine the cause of the lower-than-target values for the other three samples. | The analyst operating the instrument has been instructed to take greater care when preparing samples for analysis. We are examining the original data and are rerunning the other three samples to determine if they were mispoured as well. The examination is targeted to be completed by June 16, 2006. Further corrective action may be necessary depending on the results of the investigation. |

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Chromium, Filtered ICP labcode 0722 parmcode 1030E | Results show a marked upward shift as of May 2006. The average bias since then is 16% above the target value. | The analyst had noted the same bias and made new calibration standards in mid-July. On-line QC charts indicate that this has corrected the bias. | Line will be monitored to ensure that the new calibration standards corrected the bias. |
| Mercury, Filtered CVAF labcode 2707 parmcode 71890C | Results received as of June 15, 2006, exhibit an average negative bias of 6.5% below the target value. The most recently received data, however, show improvement. Analysis should be closely monitored until the next QADATA Summary is distributed. | Bias shifts appear to correspond with the times new calibration standards were made. Even though the stock solutions do not expire until December they appear to be degrading. | New stock solutions have arrived and new calibration standards were made. Initial indications are that the new calibration standards have corrected the bias. The analytical line will be monitored to track the effects of this change and the stock solutions will be checked monthly to monitor for degradation of the stock solutions. |
| Mercury, Whole-Water Recoverable CVAF labcode 2708 parmcode 71900D | As of early June 2006, the results show a negative bias. The average bias since then is 7.0% below the target value. | Bias shifts appear to correspond with the times new calibration standards were made. Even though the stock solutions do not expire until December they appear to be degrading. | New stock solutions have arrived and new calibration standards were made. Initial indications are that the new calibration standards have corrected the bias. The analytical line will be monitored to track the effects of this change and the stock solutions will be checked monthly to monitor for degradation of the stock solutions. |
| Nickel, Filtered ICP-MS labcode 1795 parmcode 1065G | Data received as of mid March 2006 demonstrate a positive bias. As of then, 31 out of 32 results are above the target value. No bias is evident for WWR nickel (parmcode 1067I). | This analysis has been moved from ICP-MS instrument to the new collision-cell ICP-MS instrument. New calibration standards were prepared mid-July. | The analytical line will be monitored to ensure that the move and the new calibration standards corrected the bias. To date indications are that these actions have corrected the bias. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | Since the last QADATA Summary, results for all blind samples submitted show acceptable bias and variability. No corrective action necessary. |  |  |
| Previous comments for strontium, filtered, ICP-MS, labcode 2507, parmcode 1080I |
| From QADATA Summary June 2006 | Results from the low concentration blind sample (at 31.8 ug/L) are biased on average 6.4% lower than the target value. In addition, about one-half of the results for this sample are outside of the lower control limit (-2 f-pseudosigmas). Results for all other blind samples submitted show acceptable bias and variability. | New calibration standards were prepared in May. All blind samples have been rerun using the new calibration standards and no bias is evident at this time. | Analysis will be closely monitored until next QADATA Summary is distributed. |
| From QADATA Summary April 2006 | The data are now biased less than 5% from their target value. However, three out of the last 13 results (23%) are outside -2 f-pseudosigmas. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| From QADATA Summary February 2006 | An upward trend in the data is shown starting December 2005. Data since then exhibit a positive bias of 7.1% above the target value. | There was increased scatter starting in October and the upward trend starting in December. Starting in Oct. most of the Sr results were from the new ICP-MS. | We transferred the Sr analysis to the older ICP-MS the week of January 29th and have stopped running Sr on the new instrument. There is not enough data to see the effect on the bias. If that fixes the problem we will not analyze Sr on the new instrument until we determine the cause of the problem and fix it. |
| Zinc, Whole-Water Recoverable ICP-MS labcode 2390 parmcode 1092D | The data indicate a negative bias. As of April 2006, results are on average 5.5% below the target value. The magnitude of the negative bias appears to increase as the concentration increases. | It is suspected that the higher calibration standards have higher concentrations than the target values. | New calibration standards were made in early August. The analytical line will be monitored to ensure that the new calibration standards have corrected the bias. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Filtered Electrometry labcode 2109 parmcode 29801A | Although a slight positive bias is still demonstrated, all results received as of the last QADATA Summary are within +/-1 f-pseudosigma. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Previous comments for alkalinity, filtered, electrometry, labcode 2109, parmcode 29801A |
| From QADATA Summary June 2006 | The data indicate a positive bias. Results from samples analyzed as of November 2005 are biased on average 7.8% above the target value. | Use of new instrumentation was initiated in July 2005. Prior to the new instrumentation, spike recoveries showed a negative bias. After the new instrumentation was initiated the bias became positive with decreasing magnitude and more scatter. Recoveries for reference material M-152 (target value = 6.00 mg/L) from April 2006 to present show a 0.2% positive bias. M-168 recoveries (target value = 28.2 mg/L) from March to present show a 3.4% positive bias. | At this time no corrective action will be taken. The inflection point technique for determining concentration is being investigated for use in this method. The investigation is targeted to be completed by October 1, 2006. |
| Alkalinity, Whole-Water Recoverable Electrometry labcode 0070 parmcode 90410A | Although a slight positive bias is still demonstrated, all results received as of the last QADATA Summary are within +/-1 f-pseudosigma. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Previous comments for alkalinity, whole-water recoverable, electrometry, labcode 0070, parmcode 90410A |
| From QADATA Summary June 2006 | The data indicate a positive bias. Results from samples analyzed as of November 2005 are biased on average 7.5% above the target value. | Use of new instrumentation was initiated in July 2005. Prior to the new instrumentation, spike recoveries showed a negative bias. After the new instrumentation was initiated the bias became positive with decreasing magnitude and more scatter. Recoveries for reference material M-152 (target value = 6.00 mg/L) from April 2006 to present show a 0.2% positive bias. M-168 recoveries (target value = 28.2 mg/L) from March to present show a 3.4% positive bias. | At this time no corrective action will be taken. The inflection point technique for determining concentration is being investigated for use in this method. The investigation is targeted to be completed by October 1, 2006. |
| Fluoride, Filtered ISE labcode 0031 parmcode 0950B | Although a slight positive bias is still demonstrated, most results received as of June 2006 are within +/-1 f-pseudosigma. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Previous comments for fluoride, filtered, ISE, labcode 0031, parmcode 0950B |
| From QADATA Summary June 2006 | The data exhibit a marked upward shift starting April 2006 with some improvement shown in May. Results from samples analyzed as April are biased on average 18% above the target value. | The upward shift in bias seen in the BQS charts was not evident in the online QC charts for the months of April and May. The analytical system, however, has been thoroughly cleaned several times and all pump tubing replaced. A new reference and fluoride electrode were installed on May 17. | The high bias evident in April seems to have diminished. The Nutrient Unit will closely monitor line performance and BQS charts to see what effect the change of electrodes has made. If necessary, new standards will be made. |