General procedures for CO2 and H2O calibration of infrared gas analyzers (IRGA) and secondary gas standards.

T. Ocheltree, H. Loescher

- Rationale

The purpose of this general procedure is to determine new calibration coefficients for an IRGA and to develop secondary gas standards from primary gas standards.  Secondary gases are used for both field-site and in-house IRGA calibrations.  Using primary standards for these ‘routine’ calibrations is generally unfeasible due to their high cost or loss of consistency.

B. Equipment and Setup

Needed are at least (3) primary standard gases at known concentrations.  We use the WMO-CMDL standards from NOAA with accuracy of ± 0.01 ppm (http://www.cmdl.noaa.gov/ccgg/refgases/index.html).  Other groups calibrate their primary standards using GC-MS process and dedicate a set of tanks used only to develop a second set of standards used for routine IRGA calibrations.  In this way, their set of primary standards maintains assurance and consistency across time in the calibration of the secondary standards.  WMO-CMDL standards need to be maintained as well, periodic rolling and re-calibration every 2 years are recommended.  The chosen gases need to span the range of CO2 concentrations expected in the field.

A dedicated closed-path IRGA with known history is used to calibrate the secondary standards.  IRGA is mass flow, pressure and temperature controlled.  We place the IRGA inside a temperature-controlled chamber.  Data is collected at 1 sec by a datalogger (model CR23X, Campbell Scientific, Logan, UT).  We use a configuration of valves and solenoids controlled by the datalogger to switch the flow of different standard gases.  Scrubbed N2 is used for reference gas.

The mv output signals used to convert to molar fraction of CO2 are different for each model of IRGA.  From the li-6262, the raw [CO2] and [H2O] signals are normalized by pressure, and standard CO2 concentration is normalized by temperature, such that 1) the CO2 mV output signal is multiplied by P/P0, where P is the pressure of the cell (KPa) and P0 is the average cell pressure over the entire calibration procedure (KPa), and 2). The CO2 concentration is multiplied by T0/T, where T is the cell temperature (˚C) and T0 is the average cell temperature (˚C) during the entire calibration procedure.  To determine new calibration coefficients for the IRGA, fit a 3rd-order polynomial such that, 					Eq. 1

where , , and  are constants (i.e., the calibration coefficients).  Make sure to use the mV signal as the independent variable when fitting this curve.  Note: Do not use Sigma-Plot to fit this polynomial.

The CO2 and H2O absorbance signals are used by the li-7000 to determine the calibration coefficients.  The normalization for these signals is slightly different, such that each absorbance signal is divided by mean cell pressure (KPa) and the scalar concentrations (predetermined) divided by mean cell temperature (˚K) for each concentration sampled.  Hence, to determine new calibration coefficients for the IRGA, fit a 3rd-order polynomial such that,

Eq. 2

where , , and  are constants.  Make sure to use the mV signal as the independent variable when fitting this curve.  Note: Do not use Sigma-Plot to fit this polynomial.

C.  Calibration Procedure

We first determine new calibration coefficients for our dedicated IRGA by the procedure outlined above by switching between primary standard gases across three different temperatures, 25, 30, and 35 °C (of course zero-ing the IRGA first).  After the temperature has stabilized, we then sample the gases from lowest to highest concentration, and then vise versa (i.e., ramp schedule).  In this way, any bias due to the previous gas can be removed.  After cycling through the gases at one temperature, we increase the temperature of the IRGA and sample the gases again.

We then repeat this process (with the new calibration coefficients) with both the primary standards and the secondary standards to be calibrated.  Concentrations of secondary standards are then determined using the same (new) calibration coefficients.  We also check the noise to signal ratio, as an indication of repeatability.

The procedure to calibrate the IRGA for water vapor is similar.  We use a dew point generator (model li-610, Li-cor Inc., Lincoln, NE) to automate the production of water vapor at dew points ranging 6 -20 °C (the equivalent to a vapor pressure of 0.8  -2.3 kPa) at 2 °C increments with cell temperatures set at 30 and 35 °C.  A chilled-mirror (model D2 with Optica Monitor, Eastern Instruments Inc.) is used as our primary dew point standard when calculating new coefficients.  We have found that water vapor made by the dew point generator can drift (particularly when water levels in the cooler or reservoir are low) or can produce ‘slugs’ of higher water vapor concentration as indicated by the chilled-mirror.  Hence, the chilled-mirror allows us to independently monitor the behavior of the generator, assuring data quality.  The dewpoint generator is held at each dew point temperature for 5 min before switching to the next dew point.  We filter out the first 2 min in each 5 min sampling period to allow the dewpoint generator and IRGA to stabilize at each new dew point temperature.  We then develop new calibration coefficients outlined above.

D. Field Calibration

Once new coefficients are developed, installed and tested, we re-zero and span for use in the field.  Field calibrations must be performed at the same sample cell temperature and pressure as you would expect by in situ conditions.  Often this is done at the same flow rates as used in the field.  Sometimes this is impractical because this practice results in ‘burning’ lots of secondary gas tanks if sampling at ~ 9 lpm.  An alternative methodology is to lower the flow rates but maintain the pressure and temperature of the cell (for example, 30 ˚C and at 85 KPa at 1 lpm).  Likewise, the ‘zero-ing’ should also be done with flow and pressure of the reference cell as that used in the field.

NOTE: BE SURE TO TURN OFF THE TANKS WHEN YOUR CALIBRATION IS COMPLETE OR WHEN YOU LEAVE FOR THE DAY.  THIS WILL HOPEFULLY AVOID INADVERTANTLY VENTING A TANK

Other notes: care should be used when handling the CMDL-WMO gas standards and it is advised to read the recommendations at http://www.cmdl.noaa.gov/ccgg/refgases/airstandard.html.  Particularly, any contaminations should be avoided, which includes the need for low-volume high precision regulators, periodically rolling the tanks, and having the standard tanks re-analyzed every 2 years.