##### QADATA Summary – May 2004

The Inorganic Blind Sample Project (IBSP) reviews 108 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 79 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-control samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 10/01/03 through 05/17/04; however, longer-range charts are used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the districts’ data should also affect the data displayed in the IBSP quality-control charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have been affected by a problem independent of the laboratory are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report to help provide a clearer perspective of the overall concern.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY04charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Arsenic, Filtered, GFAA, Labcode 2160, Parmcode 1000C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Filtered, ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	4

Chromium, Filtered, GFAA, Labcode 1936, Parmcode 1030I  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	4

Strontium, Filtered, ICP-MS, Labcode 2507, Parmcode 1080I . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Uranium, Filtered, ICP-MS, Labcode 1797, Parmcode 22703G. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Calcium, Filtered, ICP, Labcode 0659, Parmcode 0915D  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Sodium, Filtered, ICP, Labcode 0675, Parmcode 0930C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	6

Silica, Filtered, Colorimetric, Labcode 0056, Parmcode 0955C  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Residue on Evaporation, Filtered, Gravimetric, Labcode 0027, Parmcode 70300A . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

OWQRL

Analyte

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	9

Thallium, Whole-Water Recoverable, GFAA, Labcode 3131, Parmcode 1059(0)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	10

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Filtered GFAA labcode 2160 parmcode 1000C | Results for samples analyzed as of March 2004 demonstrate no bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL continues to closely monitor analysis. |
| Previous comments for arsenic, filtered, GFAA, labcode 2160, parmcode 1000C |
| From QADATA Summary March 2004 | This analytical line is still demonstrating a slight positive bias (6.1%). Nineteen out of 20 values are greater than their median value. | The slight positive bias is also evident on QC charts for third-party check samples (TPC). New standards were made and maintenance was performed during the week of March 8th, 2004. | After making new standards and performing maintenance, the positive bias is no longer evident. This line will continue to be monitored for bias. |
| From QADATA Summary January 2004 | Data since October 2003 indicate a slight positive bias (7%). Since October there are 11 values (all results) greater than their median value. | New standards were made in November of 2003 including Continuing Calibration Verifications (CCV) and Third-Party Check standards (TPC). Both the CCV’s and TPC’s for arsenic are showing a positive bias on QC charts. | New calibration standards will be remade and verified as of February 6th, 2004. The element will continue to be monitored closely for a bias. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | As of March 2004, results indicate the positive bias is less than 5%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL continues to closely monitor analysis. |
| Previous comments for cadmium, filtered, ICP-MS, labcode 1788, parmcode 1025G |
| From QADATA Summary March 2004 | Data from January 2004 through March exhibit a slight positive bias (6%). Almost all results are greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | Due to the on-going monitoring of internal QC, a positive bias is still not evident on NWQL QC charts. To verify this, an additional TPC from a different source than is routinely used by ICP-MS was run during the week of March 8th, 2004, to verify the ICP-MS calibration curve. Calibration and QC are acceptable. | The issue of possible interferences that could affect cadmium is being discussed with the NWQL’s Method Research and Development Program. More research needs to be done. Also, internal QC and IBSP charts will continue to be monitored for bias and variability. |
| From QADATA Summary January 2004 | Data for the last six months (June 2003 through December) demonstrate a slight positive bias (6%) with all 48 values greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of December 2003. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | Data received as of March 2004 indicate a negative bias (8.5%) . | The bias is not evident on NWQL QC charts. Calibration standards have been checked for accuracy and have been verified. Study is being conducted to establish if any instrument modifications should be performed. | NWQL continues to closely monitor analysis for bias and evaluate the need for instrument modifications. |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030I | Data received since the last QADATA Summary indicate no bias, and all values are within +/-2 f-pseudosigma. No corrective action necessary. |  | NWQL continues to closely monitor analysis. |
| Previous comments for chromium, filtered, GFAA, labcode 1936, parmcode 1030I |
| From QADATA Summary March 2004 | Since late-January 2004, variability has increased and a negative bias has (13.5%) developed. | This line had three samples that fell outside (low) of the +/-2 f-pseudosigma which contributed to the large negative bias. One, with a target value of 6.12, was reported at 1.45. The other two values were between 2 and 3 f-pseudosigma. The negative bias on BQS control charts is not evident on NWQL QC charts. In fact, a slightly high bias was evident on NWQL QC charts at the end of February. | Maintenance was performed on the instrument during the week of March 8, 2004, and early evidence suggests that the positive bias has decreased. Also, the three samples mentioned in the comments section were rerun and were within acceptable limits. The recalculated bias is 4.5%. Analyst will continue to monitor bias and variability and perform maintenance as needed to alleviate both of these problems. NOTE: Three environmental samples associated with the set that had the low value of 1.45 were updated in LIMS and released to the customer. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | As of March 2004 variability has increased. Four out of 16 values (25%) are outside of control limits (+/-2 f-pseudosigma) as of March 2004. Data as of early-April suggest the problem has been corrected. No corrective action necessary. |  | NWQL continues to closely monitor analysis. |
| Uranium, Filtered ICP-MS labcode 1797 parmcode 22703G | As of March 2004 variability has increased. Four out of 15 values (27%) are at or outside of control limits (+/-2 f-pseudosigma) as of March 2004. | NWQL QC charts show increased variability. | New uranium standards were purchased, prepared, and verified by April 19th. Since then, NWQL QC Charts show reduced variability. The NWQL will continue to closely monitor the analysis for variability. |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Calcium, Filtered ICP labcode 0659 parmcode 0915D Magnesium, Filtered ICP labcode 0663 parmcode 0925C Sodium, Filtered ICP labcode 0675 parmcode 0930C | Data received since the last QADATA Summary indicate a positive bias. In addition, more values at or outside of +/-2 f-pseudosigma have been received. Results from the new instruments (put into service June 2003) still indicate high variability relative to results from the old instruments. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last QADATA Summary (March 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. Tests have been performed that rule out non-instrument related contamination. Tests are being conducted to determine whether instrument contamination and carry-over is present. The positive bias for calcium and magnesium on BQS control charts is also evident on NWQL QC charts, prompting the preparation of new standards. The positive bias for sodium is not evident on NWQL control charts. | NWQL continues to investigate the cause of the outliers on runs with acceptable QC by running tests on the instrument and determining whether contamination and carry-over is present. New standards for calcium, magnesium, and sodium have been prepared and were validated on May 12th, 2004. NWQL will continue to closely monitor the analysis for bias. |
| Previous comments for calcium, magnesium, and sodium by ICP (next page) |
| From QADATA Summary March 2004 | New instrumentation was implemented in June 2003. A comparison of results from the old instruments (TJA) to results from the new instruments (PE) indicates an increase in variability. The OWQRL is not demonstrating variability problems for these constituents. | The increase in variability is due in part to several outliers, some of which have been due to human error. Since the last data summary (January 2004), however, all outliers have been on runs with acceptable QC. Their cause is not known at this time. | As stated in the January 2004 QADATA Summary for cadmium WWR GFAA (labcode 1555), use of an additional QC check sheet that is included in all data packets was implemented to help alleviate incorrect data going in LIMS. This practice was implemented for ICP processes in October 2003, and data shows that it has been useful in eliminating some errors. We will continue to investigate the cause of the outliers on runs with acceptable QC. |
| Silica, Filtered Colorimetric labcode 0056 parmcode 00955C | From mid-November 2003 through December, results for two reference materials exhibit a large positive bias. Results received as of January 2004 show no overall bias, and all values are within +/-2 f-pseudosigma. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL continues to closely monitor analysis. |
| Residue on Evaporation, Filtered Gravimetric labcode 0027 parmcode 70300A | Data received since the last QADATA Summary indicate no bias, and all values are within +/-2 f-pseudosigma. No corrective action necessary. |  |  |
| Previous comments for residue on evaporation, filtered, gravimetric, labcode 0027, parmcode 70300A |
| From QADATA Summary March 2004 | A slight negative bias (5%) has developed starting December 2003. | Between December 2003 and February 2004 the NWQL also noticed a low bias with our on-line QC for M-162. The blank and M-152 on-line QC also show a very slight low bias; however, the bias is not as evident. The bias went away near the beginning of March 2004, according to the M-162 on-line QC. All on-line QC are around the expected means as of March 2004. During this time the analyst reported problems with the balance weight being sent to the computer in a consistent manner. The balance doors were adjusted and the timing routine was modified. | The balance communication will continue to be monitored using on-line QC for a reoccurrence of the low bias |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | The most recently received data indicate that this method is still demonstrating variability beyond acceptable levels. Of the last seven values received, one has been within the control limits. | Due to the pending closure of the laboratory, new standards were not ordered . | Current stock was used for new standards. |
| Previous comments for arsenic, whole-water recoverable, GFAA, labcode 3119, parmcode 1002(1) |
| From QADATA Summary March 2004 | Since October 2002, data for this analytical line have demonstrated variability beyond acceptable levels. Since October 2002, nine out of 28 points (32%) are outside of +/-2 f-pseudosigma with three of 28 (11%) outside of +/-3 f-pseudosigma. The most recently received data indicate continuing problems with the method variability. IBSP recommends that the source of the problem be determined before releasing any more data from environmental samples. | The most recent problems seem to come from the standards used for arsenic. The standard mixes used for ICP have been used for GFAA analyses instead of the standards normally used for GFAA. The arsenic standard was also found not to be matrix matched for the digestion. | New standards for the GFAA will be ordered, checked and matrix matched for whole-water recoverable samples. Any samples requesting arsenic that we have will then be rerun. |
| From QADATA Summary January 2004 | The data still indicate high variability. Two of the last seven values received (29%) are outside of +/-2 f-pseudosigma. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| From QADATA Summary Nov 2003 | High variability is still demonstrated. Three of the last seven values received (43%) are outside of +/-2 f-pseudosigma. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| From QADATA Summary Sept 2003 | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 f-pseudosigma with 3 of 19 (16%) outside of +/-3 f- pseudosigma. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 f-pseudosigma with 3 of 13 (23%) outside of +/-3 f-pseudosigma. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 f-pseudosigma. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |
| Thallium, Whole-Water Recoverable GFAA labcode 3131 parmcode 1059(0) | As of January 2004 results indicate a negative bias (22%). This bias is not evident for other methods analyzing for thallium. | Problems with Tl standards. Due to pending closure of the laboratory, new standards were not ordered. | Current stock was used for new standards. |