DOE/SC-ARM/TR-062

Shortwave Spectroradiometer (SWS) Handbook

May 2007

Peter Pilewskie

John Pommier

Work supported by the U.S. Department of Energy,

Office of Science, Office of Biological and Environmental Research

Contents

1	General Overview	1

2	Contacts	2

3	Deployment Locations and History	2

4	Near-Real-Time Data Plots	2

5	Data Description and Examples	3

6	Data Quality	5

7	Instrument Details	6

Figures

1	SWS and Monitor in the Darkroom	1

2	Schematic for the Connection of the SWS.	7

# General Overview

The Shortwave Spectroradiometer (SWS), measures the absolute visible and near infrared spectral radiance (units of watts per meter square per nanometer per steradian) of the zenith directly above the instrument.  The SWS is a moderate resolution sensor comprised of two Zeiss spectrometers (MMS 1 NIR enhanced and NIR-PGS 2.2) for visible and near-infrared detection in the wavelength range 350 – 2170 nm.  The sampling frequency is 1 Hz.  The spectral resolution is 8 nm for the MMS 1 NIR and 12 nm for the NIR-PGS 2.2.  The light collector is a narrow field of view (1.4() collimator at the front end of a high-grade custom-made fiber optic bundle.  The SWS does daily daytime measurements.  The SWS is located in a darkroom, constructed by Southern Great Plains (SGP) site personnel within the optical trailer, to permit calibrations to be conducted without the necessity of moving the instrument to a different location.  Calibrations are performed at regularly scheduled times using the ARM 12” integrating sphere.  The SWS measurements can be used to:

- Retrieve cloud optical depth, particle size and cloud water path.

- Test the cloud optical depth retrieval for overcast and broken cloud fields.

- Validation/comparison with SGP surface remote sensors and future cloud intensive operational period campaigns.

- Multivariate analysis to derive information content in hyper spectral data sets and to improve cloud retrieval algorithm development.

- Compare with radiative transfer models for testing and validating retrieval procedures.

Figure 1.  SWS and monitor in the darkroom.  Insert shows the light collector located at the top of the chimney.

Sites/Facilities

Southern Great Plains (http://www.arm.gov/sites/sgp.stm)

Central Facility (http://www.arm.gov/sites/sgp/central.stm)

# Contacts

## Mentor

Scott Kittelman

University of Colorado

Department of Atmospheric and Oceanic Sciences

Folsom Stadium, Gate 7, Rm 255

UCB 311

Boulder, CO 80309-0311

(303) 492-4248

E-mail:  alan.kittelman@colorado.edu

Patrick McBride

Department of Atmospheric and Oceanic Sciences

Duane Physics, Rm D-335

UCB 311

University of Colorado

Boulder, CO 80309-0311

(303) 492-8653

E-mail:  mcbridep@colorado.edu

## Instrument Developer

Warren Gore

NASA Ames Research Center

MS 245-4

Moffett Field, CA 94035

(650) 604-5533

# Deployment Locations and History

| Serial Number | Property Number | Location | Date Installed | Date Removed | Status |
| --- | --- | --- | --- | --- | --- |
| ? | ? | SGP/CF | April 2006 | N/A | Operational |

Data collected is available from the Atmospheric Radiation Measurement (ARM) Program’s archive (http://www.archive.arm.gov).

# Near-Real-Time Data Plots

Will be available at the ARM web site:  Data plots and other data products (http://www.arm.gov/data/data_plots.stm)

# Data Description and Examples

## Data File Contents

The SWS data files are stored in both raw and calibrated files at the ARM archive.  File names follow the ARM naming convention for the raw data files as follows:

sgpswscf.00.yyyymmdd.hhmmss.dat.  Where yyyy ( year (i.e. 2006), mm ( month, dd ( day, hh ( hour, mm ( minutes and ss ( seconds.  The raw data files will be calibrated (units of Wmˉ²nmˉ¹srˉ¹) and converted into netcdf format by ARM at the time of ingest.

### Primary Variables and Expected Uncertainty

Absolute spectral radiance of the zenith above the instrument in units of Wmˉ²nmˉ¹srˉ¹.

#### Definition of Uncertainty

Prior to deployment at the SGP, the SWS was calibrated for absolute spectral power at the Airborne Science and Technology Laboratory at NASA Ames Research Center using the 30” integrating sphere, which has been calibrated according to NIST standards, and a LabSphere 12” integrating sphere.  The 12” sphere is used at the SGP site to calibrate the SWS.

The absolute accuracy of the spectral radiance for the SWS relies upon the accuracy of the transfer standard from the 30” sphere to the 12” sphere.  The 30” sphere has an error over the spectral range, 300 – 2200 nm, of between 1 – 2%.

### Secondary/Underlying Variables

This section is not applicable to this instrument.

### Diagnostic Variables

These are variables that monitor the health of the SWS.  The first three are concerned with the temperature of the two spectrometers and the box in which they are encased, while the latter variables monitor voltages and indicate if a critical component is failing.

| Channel | Quantity Measured | Unit |
| --- | --- | --- |
| 0 | Si spectrometer. Normal temperature is 27°C | C |
| 1 | InGaAs spectrometer temperature. Normal temperature is -10°C. | C |
| 2 | Inside box temperature. This temperature reflects the ambient temperature which should be around 22°C. | C |
| 3 | PC/104 + 5VDC (nominal value = +5V) | V |
| 4 | PC/104 + 12VDC (input voltage*3, nominal voltage = +4V) | V |
| 5 | PC/104 – 12VDC (input voltage*3, nominal voltage = -4V) | V |
| 6 | PS2 +5VDC (nominal voltage = +5V) | V |
| 7 | PS2 +12VDC (input voltage*3, nominal voltage = +4V) | V |

### Data Quality Flags

Not available for this instrument at this time.

### Dimension Variables

All SWS data files contain Universal Time Coordinates (UTC) and the date.

## Annotated Examples

This figure is of sample spectra collected on 1 May 2006 during the period 1500 – 1900 UTC (10:00 – 14:00 local) time.  The spectra are not averaged; they are the 1 Hz spectrum collected at the indicated times.

The above plot is a radiance time series at nine SWS wavelengths.  The times are from 1330 – 2200 UTC.  The small gaps on the hour (1400, 1500 etc.) are periods when the shutter was closed to obtain the dark signal.  There was clear sky until around 1530 UTC when first scattered clouds then a thicker (1900 UTC) cloud mass moved over head.

## User Notes and Known Problems

This section is not applicable to this instrument.

## Frequently Asked Questions

There are no FAQ’s at the present time.

# Data Quality

## Data Quality Health and Status

Data Quality for the SWS will be available from the following web site maintained by DQ Hands:  http://dq.arm.gov/

## Data Reviews by Instrument Mentor

The monthly review by the mentor is available from the ARM web site http://dq.arm.gov/

## Data Assessments by Site Scientist/Data Quality Office

Not applicable to this instrument.

## Value-Added Procedures and Quality Measurement Experiments

To be determined.

# Instrument Details

## Detailed Description

### List of Components

- Narrow field-of-view (1.4() collimator at the front end of a high grade, ultra-low OH fiber optic cable.

- Spectrometer:

Zeiss MMS 1 NIR enhanced (300-1100 nm), flat-field, 366 1/mm grating, Hamamatsu Si 256-element linear diode array.  Operating temperature regulated to 27(C by heater.

Zeiss NIR-PGS 2.2 (900-2200 nm), 300 1/mm flat grating, Hamamatsu InGaAs 256-element linear diode array.  Operating temperature regulated to -10(C by thermoelectric cooler.

- Automated shutter for dark current measurement.

- A 933 MHz Pentium III based computer (PC104 format) for automated data acquisition and control.

- A USB interface between the computer and the spectrometers.

- Operating system is Linux Fedora.

- Internal data storage on 1 GB compact flash card; remote data collection via Ethernet interface.

### System Configuration and Measurement Methods

The SWS is located in a dark room, specially constructed by SGP site personnel, in the south east corner of the Optical Trailer at the SGP Central Facility.  Placing the SWS in the dark room obviates the need to move the SWS to perform calibrations.  The light collector is attached to a high quality fiber optic which is passed through the port at the top of the optical trailer to the SWS.

Figure 2.  Schematic for the Connection of the SWS

### Specifications

Wavelengths Measured:  256 channels in the Si (300 – 1100 nm), 256 channels for the InGaAs (900 – 2200 nm).

Instrument Field of View:  The field of view of the instrument is 1.4(.

Sampling Interval:  Measurements are made at the rate of one per second during the day.  The SWS does not collect data between the hours of 0200 – 1100 UTC.

Integration Time:  The integration times vary during the course of the year but range from 75 – 100 ms for the Si detector and range approximately 150 – 250 ms for the InGaAs detector.

## Theory of Operation

Photons incident on the collimator at the fore-optics of the light collector travel through the multi-optical fiber to the SWS where the fiber is then bifurcated with the light passing through entrance slits to either the Si grating or the InGaAs grating.  There the wavelength components are separated then imaged onto their respective diode array.  The array is then read by an electronic interface which passes the data to the computer via a USB connection.  Dark signals are obtained each hour, at the same integration time that was used to measure the intensity, by closing the shutter located at the base of the light collector.

## Calibration

### Theory

The SWS is calibrated annually at the NASA Ames Airborne Science and Application Laboratory using a 30” integrating sphere.  The 30” sphere has been calibrated according to NIST standards.  The calibration from the 30” sphere is transferred annually to the ARM Southern Great Plains on-site 12” integrating sphere.  The ARM 12” sphere is used weekly at the SGP site to monitor for any changes or degradation in the SWS response.  If necessary, the calibration obtained with the ARM 12” integrating sphere is applied to adjust the primary SWS response function obtained annually at NASA Ames.

### Procedures

SGP site personnel built a darkroom to both house the SWS instrument and the ARM 12” integrating sphere used for weekly on-site calibrations.  This allows the SWS to be calibrated without having to move the instrument to another part of the site.  To perform a calibration, the light collector is lowered through the port in the ceiling and placed in a holder opposite the 12” integrating sphere.  The holder is aligned with the central axis of the exit port on the 12” sphere to insure the light collector is situated in the identical location each time it is calibrated.  A dark signal is taken before each calibration.  Following the calibration the light collector is then returned through the port to the top of the Optical Trailer.

### History

No history is available at this time.

## Operation and Maintenance

### User Manual

A user manual has been prepared and provided to the SGP site personnel.  An electronic version is available upon request from the mentor in either PDF or Postscript format.

### Routine and Corrective Maintenance Documentation

Instrument preventative maintenance reports are can be accessed at the SGP:  http://www.ops.sgp.arm.gov/

### Software Documentation

Software documentation is found in the User Manual (see 7.4.1 above).

### Additional Documentation

None available for this instrument.

## Glossary

See the ARM Glossary (http://www.arm.gov/about/glossary.stm)

## Acronyms

See the ARM Acronyms and Abbreviations.  (http://www.arm.gov/about/acronyms.stm)