STATION RECORD SHEET

# SYNOP or CLIMATE/SYNOP RECORD SHEET

Station Name:	Keetmanshoop

Date Station Established:  1948-10-01 as a rainfall station.	Synop No: 68312

Digitized records from May 1970, that is Temperature and other parameters.

Latitude:	26deg 32min SOUTH		Longitude: 18deg 06min EAST

Height above MSL:	1066m			Point used: Met Enclosure

Observer: J.Jacobs

Phone No: 063 223664

Postal Address:				Address for policy matters:

P O .Box 10          				Namibian Meteorological Service

Keetmanshoop					P/Bag 13224

Namibia            				Windhoek

Namibia.

# Location Indicator:

The station is located in the Keetmanshoop Air Port.

# Remarks:

## Clear site and instruments are well exposed

# Instructions for locating station by road:

## The station is within the airport enclosure.

Synop reporting times: 0300z, 0600Z 0900z, 1200Z, 1500z and 1800z; everyday’ except that on Saturdays the 1500z observations are not taken..

# Metar reporting times: Metars are reported hourly, however the 1600 and 1700hrs observations are skipped. The metar reports on Saturdays are not taken at 1300hrs, 1400hrs, 1500hrs and 1700hrs.

# Transmission of reports: All reports are sent to Windhoek by fax and cell phone.

Inspections by: D. F. Molotsi and E. Magagwa		Date: 04/10/06

Returns furnished: Data in the diaries (log books) are transcribed on to monthly climate return forms and sent to the National Meteorological Service Head quarters in Windhoek.

.

# Annual supplies: Thermometers, pens, ink and wicks.

# Non- Instrumental (books, charts, scales etc):. All charts, and daily log books

## MANUAL RAINGAUGE:

Type:						5 inch copper standard rain-gauge

Purpose:				 	climate

Inst Date:					not known

Height of rim above MSL:			121cm

Glass measure type and number:		20mm measuring glass and in good condition

Location/Remarks:				In the Met enclosure.

### AUTOMTATIC RAINGAUGE

Type:						Daily

No:						W5364

Chart No:					No number locally made charts

Height of rim above ground level:		124cm

Clock:					           Daily spring clockwork mechanism

Drum:

Location/Remarks:				Met enclosure and well exposed.

## THERMOMETER SCREEN:

Type:						Stevenson screen

Support:					metal stand

Inst date:					not known

Location/Remarks:				Installed in the Met enclosure. The screen is worn out the door is completely broken. Need to be replaced urgently.

THERMOMETERS:	TYPE: 	MAKE & S/ NO:	           CERTIFICATE

Dry bulb			sheathed	Lambrecht/ 5325                               no

Wet bulb			sheathed	Therm Scheider/ 852389                   no

Maximum			sheathed	CW. Price & Co/ 2238.05                 no.

Minimum			sheathed	CW. Price & Co/ 894/91                   no.

Grass minimum		n/a

5cm earth			n/a

10cm earth			n/a

20cm earth			n/a

30cm earth			n/a

100cm earth			n/a

120cm earth					Zeal 2394                                         no

Remarks: The earth thermometers are not available except the 120cm which is also broken.

## REMOTE INDICATING ISTRUMENTS

Type:

Location of sensor:

Location of indicator:

The are no remote indicating instruments.

## BAROMETER

Type:					                       PA21 & FWilh Lambrecht

# No:                                                                             26550

Installed:					            Records are not available.

Location:						Observatory

# New or old convention:                                            Both old

# Attached thermometer in Degrees:			Degrees/Kelvin Scale

# Height of cistern above MSL:			87.5cm

# Gravity of station:

# When and how height obtained:			10/10/06 - by the use of tape measure

# Correction cards for:

# Remarks: The PA21 is mostly used and the mercury barometer is used in when they is power blackout for comparison.

## BAROGRAPH

Type:							Micro Barograph

Maker:						Thies Clima

No:							3.0811.10.000

Installed:						Date is not known.

Clock:							Daily Spring Clockwork Mechanism.

Chart no:         					Met 3/2/1/3

# Remarks: The instrument is in a good working condition.

# .

## SUNSHINE RECORDER

Height above ground: 	136cm

Installed:			Date of installation is not known.

# Location/mounting/remarks: The instrument is mounted on a still base in the met enclosure. It is well maintained.

## RADIATION RECORDER

# Sensor Make and Type:

# Sensor Make and Type:

# Sensor Make and Type:

# Location of sensors:

# Recorder Type:

# Recorder location:

# Solar Panel No:

# Remarks: The instrument is not available at this station.

## HORIZON PROFILE

# Date prepared: The date of preparation is not known.

# Observation Point: The reference point are hills about 20km south of the station

# Remarks:

## EVOPORIMETER

Type:							Class A evaporation pan

Gauge Type:						ruler scale

# Dipstick No:                                                              not available

Height of Rim above ground level:			19cm

Location/Remarks:					Met enclosure and the pan is kept clean

## THERMOGRAPH

Type:

Maker:

Installed Date:

# Clock:

Chart No:

This is instrument is not available at this station.

## EARTH THERMOGRAPH

Type:	This instrument is not available.					.

# Maker:

# Installed Date:

# Clock:

# Chart No:

# Location of Sensors:

# Remarks: This instrument is not used at this station.

## HYGROGRAPH

Type:

Maker:

Installed Date:

# Clock:

## Chart No:

Remarks:

This instrument is not available at this station.

## CLOUD SEARCHLIGHT

Type:

# Baseline:

# Location of Searchlight:

# Stand Construction:

# Observation Point:

# Alidade Type:

# Switch Type:

# Remarks: This instrument is not available at this station.

## CUP COUNTER ANEMOMOTER

Type:

Make:									Vaisala

Date Installed:							not known

# Can station take only an Angled faced unit

Height of head above ground:					2m

Location:								Met square enclosure

## ANEMOMOTER INSTRUMENTS

# Cup Counter:

# Cup Contact:

# Cup Generator:

# Wind Vane:

# Recorder, speed:

# Recorder, combined:

# Indicators Located:

# Recorder Type:

# Chart Drive:

# Recorder Location:

# Resistance Box Location:

# Power Transformer Location:

# Power point location:

# Remarks:

## MAST

Type:

Height of Head above Ground Level:

# Boom orientated

Remarks, There is no conventional wind system at this station..

## ADDITIONAL INSTRUMENTS:

There are no additional instruments at this station.

# General Exposure of Station:

The instruments at this station are well exposed. There are no obstructions around the station.

# Exposure of the Instrument enclosure:

The instruments are well exposed in a wide open meteorological enclosure.

Local Effects: Dust storms occur once in a while.