Title: V9.1 SAS System Output
Author: SAS Version 9.1
Last saved by: dbeaty

09:03  Monday, November 13, 2006  

FY06 LT-MDL Review – Whole Water Recoverable, ICP

09:03  Monday, November 13, 2006  

LTMDL Review for Calcium, Water, Unfiltered, ICP

09:03  Monday, November 13, 2006

09:03  Monday, November 13, 2006  

LTMDL Review for Magnesium, Water, Unfiltered, ICP

09:03  Monday, November 13, 2006  

LTMDL Review for Chromium, Water, Unfiltered, ICP

09:03  Monday, November 13, 2006  

LTMDL Review for Iron, Water, Unfiltered, ICP

09:03  Monday, November 13, 2006  

LTMDL Review for Manganese, Water, Unfiltered, ICP

Analyte | TestID | Lab Code | Calculated Spike LT-MDL | Calculated Blank LT-MDL | Blank Corrected LT-MDL | LT-MDL in use | Update LT-MDL To: | Update LRL To: | Update?
Calcium | 00916C | 2356 | 0.00361 | 0.00686 | 0.0064 | 0.02 | 0.007 | 0.014 | Y
Magnesium | 00927C | 2362 | 0.00123 | 0.00104 | 0.0004 | 0.004 | 0.001 | 0.002 | Y
Chromium | 01034G | 2357 | 0.39118 | 0.37063 | 0.2291 | 1 | 0.4 | 0.8 | Y
Iron | 01045C | 2359 | 3.40572 | 2.02198 | 1.926 | 3 | | | N
Manganese | 01055D | 2363 | 0.16789 | 0.11662 | 0.0403 | 0.3 | 0.2 | 0.4 | Y

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
00916C | 2356 | 047001W | 32 | 0.07 | 0 | 0.9422 | Normal
| | 04Bk | 25 | . | 0 | 0.0000 | Non-normal
| | 057001W | 30 | 0.07 | 0 | 0.9827 | Normal
| | 057008W | 24 | 0.07 | 0 | 0.9987 | Normal
| | 05Bk | 30 | . | 0 | 0.0000 | Non-normal
| | 066103W | 14 | 0.035 | 0 | 0.5162 | Normal
| | 06Bk | 10 | . | 0 | 0.8055 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
00927C | 2362 | 047001W | 32 | 0.014 | 0 | 0.0000 | Non-normal
| | 04Bk | 25 | . | 0 | 0.0038 | Non-normal
| | 057001W | 30 | 0.014 | 0 | 0.0000 | Non-normal
| | 057008W | 23 | 0.014 | 0 | 0.9421 | Normal
| | 05Bk | 30 | . | 0 | 0.0000 | Non-normal
| | 066103W | 14 | 0.006 | 0 | 0.2297 | Normal
| | 06Bk | 9 | . | 0 | 0.1473 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
01034G | 2357 | 047001W | 32 | 1.75 | 0 | 0.0888 | Normal
| | 04Bk | 26 | . | 0 | 0.1654 | Normal
| | 057001W | 31 | 1.75 | 0 | 0.1895 | Normal
| | 057008W | 23 | 3.5 | 0 | 0.7579 | Normal
| | 05Bk | 31 | . | 0 | 0.0632 | Normal
| | 066103W | 14 | 1.3 | 0 | 0.4225 | Normal
| | 06Bk | 10 | . | 0 | 0.5872 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
01045C | 2359 | 047001W | 32 | 15.75 | 0 | 0.0002 | Non-normal
| | 04Bk | 26 | . | 0 | 0.2641 | Normal
| | 057001W | 31 | 15.75 | 0 | 0.1958 | Normal
| | 057008W | 24 | 10.5 | 0 | 0.1128 | Normal
| | 05Bk | 31 | . | 0 | 0.2789 | Normal
| | 066103W | 14 | 7 | 0 | 0.9842 | Normal
| | 06Bk | 9 | . | 0 | 0.5163 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
01055D | 2363 | 047001W | 32 | 2.1 | 0 | 0.0010 | Non-normal
| | 04Bk | 25 | . | 0 | 0.3304 | Normal
| | 057001W | 30 | 2.1 | 0 | 0.5038 | Normal
| | 057008W | 23 | 1.05 | 0 | 0.0905 | Normal
| | 05Bk | 30 | . | 0 | 0.2408 | Normal
| | 066103W | 14 | 0.2 | 0 | 0.3085 | Normal
| | 06Bk | 10 | . | 0 | 0.8790 | Normal