##### QADATA Summary – February 2006

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 100 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 07/01/05 through 02/02/06; however, longer-range charts may be used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the laboratory are removed from the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and their impact on the new data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/FY05_06charts.htm.  Laboratory and district personnel can also make their own custom QADATA retrieval to parallel their objectives by using the QADATA Application online at  http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Nutrients

Phosphorus, Filtered,  (Low-level), Colorimetry, Labcode 2331, Parmcode 0666G	4

Phosphorus, Whole-Water Recoverable,  (Low-level), Colorimetry, Labcode 2333, Parmcode 0665G	5

Trace Metals

Mercury, Filtered, CVAF, Labcode 2707, Parmcode 71890C	6

Mercury, Whole-Water Recoverable, CVAF, Labcode 2708, Parmcode 71900D	6

Strontium, Filtered,  ICP-MS, Labcode 2507, Parmcode 1080I	6

Zinc, Filtered,  ICP-MS, Labcode 1798, Parmcode 1090G	7

Zinc, Whole-Water Recoverable, ICP-MS, Labcode 2390, Parmcode 1092D	7

Majors

Alkalinity, Filtered, Electrometry, Labcode 2109, Parmcode 29801A	8

Alkalinity, Whole-Water Recoverable, Electrometry, Labcode 0070, Parmcode 90410A	8

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Mercury, filtered, CVAF, parmcode 71890C – Positive bias starting mid-August 2005.  Recent data show 6.9% bias.

- Phosphorus, filtered, colorimetry, low-level, parmcode 0666G – Positive bias starting September 2005.  Recent data show 7.6% bias.

- Phosphorus, whole-water recoverable, colorimetry, low-level, parmcode 0665G – Positive bias starting mid-July 2005.  Recent data show 6.4% bias.

- Strontium, filtered, ICP-MS, parmcode 1080I – Upward trend in the data starting December 2005.  Recent data show 7.1% bias.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Phosphorus, Filtered Low-level Colorimetry labcode 2331 parmcode 0666G | The positive bias is still present with results about 7.6% above the target value. However, the most recently received data show improvement. Analysis should be closely monitored until the next QADATA Summary is distributed. | New stock standards from a different vendor seem to have fixed the problem. | We will continue to monitor this line for bias. |
| Previous comments for phosphorus, filtered, low-level, colorimetry, labcode 2331, parmcode 0666G |
| From QADATA Summary December 2005 | All results since September 2005 exhibit a positive bias. The average bias is 5.6% above the target value. This bias is not demonstrated for phosphorus determinations by the regular level methods. | New SPEX standards were put on-line on 10/27/05. The system was rebuilt on 10/20/05. A comparison of calibration standards with the alkaline persulfate line (made from an ERA stock solution) indicates that the low-level standards may be low. However, new SRWS samples prepared on 11/08/05 show values that are very near or slightly below the MPV using the existing calibrants. New batches of ammonium persulfate reagent were prepared 3 times in November. The acidic molybdate reagent is supposed to be good for a year but will be remade and tested to see if that reagent has deteriorated. | A new batch of acidic molybdate reagent was made on 12/07/05 to test the response. If necessary, we will prepare calibration standards from the same ERA stock solution used to prepare alkaline persulfate phosphorus standards by the end of December. |
| Phosphorus, Whole-Water Recoverable Low-level Colorimetry labcode 2333 parmcode 0665G | The positive bias is still present with results about 6.4% above the target value. However, the most recently received data show improvement. Analysis should be closely monitored until the next QADATA Summary is distributed. | New stock standards from a different vendor seem to have fixed the problem. | We will continue to monitor this line for bias. |
| Previous comments for phosphorus, whole-water recoverable, low-level, colorimetry, labcode 2333, parmcode 0665G |
| From QADATA Summary December 2005 | All results since mid-July 2005 show a positive bias. Results since mid-August 2005 exhibit an average positive bias of 6.1% above the target value. This bias is not demonstrated for phosphorus determinations by the regular level methods. | New SPEX calibration standards were approved on 11/15/05 and 11/29/05. The system was rebuilt on 10/20/05. New SRWS standards were put on-line on 11/17/05 and show values very near or slightly below the MPV. A comparison of calibration standards with the alkaline persulfate line (made from an ERA stock solution) indicates that the low-level standards may be low. New batches of ammonium persulfate reagent were prepared 3 times in November. The acidic molybdate reagent is supposed to be good for a year but will be remade and tested to see if that reagent has deteriorated. | A new batch of acidic molybdate reagent was made on 12/07/05 to test the response. If necessary, we will prepare calibration standards from the same ERA stock solution used to prepare alkaline persulfate phosphorus standards by the end of December. |
| From QADATA Summary October 2005 | All results received since the last QADATA Summary are above the target value. Starting mid-August 2005, results are, on the average, more than 5% above the target value. | The NWQL agrees with the observation. | New stock standards have been ordered so new working standards can be prepared and put online by October 30, 2005. |
| From QADATA Summary June 2005 | All results received since April 2005 are higher than the target value and exhibit an overall positive bias of 5.8% above the target value. A bias of this magnitude is not evident for phosphorus, filtered (lc 2331). | Working standards prepared and put online April 4th appear to have been prepared from a degrading stock. After high bias was noted by analyst, a new stock standard was ordered. New working standards put online on May 18. | Since the preparation of new standards, the bias seems to have improved. We will continue to closely monitor the line. |

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Mercury, Filtered CVAF labcode 2707 parmcode 71890C | The data still indicate a positive bias. Results for samples analyzed as of mid-December 2005 are biased on average 6.9% higher than the target value. | Filtered mercury (LC 2707) shows a positive bias for concentrations near 0.04 ug/L. Online QC samples with higher concentrations do not exhibit this bias. | Since the bias is positive and appears to be at the low end of the calibration we will remake all reagents the week of February 12th and look for sources of possible contamination. We will continue to monitor this line for bias. |
| Previous comments for mercury, filtered, CVAF, labcode 2707, parmcode 71890C |
| From QADATA Summary December 2005 | Data indicate a positive bias. Results for samples analyzed as of mid-August 2005 are biased on average 5.5% higher than the target value. | Positive bias was noted on NWQL QC and BQS charts in November. New calibrations standards were prepared on November 28th, 2005. Most recent data indicate bias is improved. | We will continue to monitor the mercury line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Mercury, Whole Water Recoverable CVAF labcode 2708 parmcode 71900D | The previously noted positive bias is now less than 5%. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | We will continue to monitor this line for bias. |
| Previous comments for mercury, whole-water recoverable, CVAF, labcode 2708, parmcode 71900D |
| From QADATA Summary December 2005 | Data exhibit a positive bias. Results for samples analyzed as of late August 2005 are biased on average 8.4% higher than the target value. | Positive bias was noted on NWQL QC and BQS charts in November. New calibrations standards were prepared on November 28th, 2005. Most recent data indicate bias is improved. | We will continue to monitor the mercury line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | An upward trend in the data is shown starting December 2005. Data since then exhibit a positive bias of 7.1% above the target value. | There was increased scatter starting in October and the upward trend starting in December. Starting in Oct. most of the Sr results were from the new ICP-MS. | We transferred the Sr analysis to the older ICP-MS the week of January 29th and have stopped running Sr on the new instrument. There is not enough data to see the effect on the bias. If that fixes the problem we will not analyze Sr on the new instrument until we determine the cause of the problem and fix it. |
| Zinc, Filtered ICP-MS labcode 1798 parmcode 1090G | This method has been discontinued by the NWQL. The replacement method is zinc, filtered, cICP-MS (collision cell), labcode 3138, parmcode 1090I. |  |  |
| Previous comments for zinc, filtered, ICP-MS, labcode 1798, parmcode 1090G |
| From QADATA Summary December 2005 | Data still exhibit a negative bias. Results from samples analyzed late August 2005 and after are biased on average 9% below the target value. This bias is not evident for zinc determinations by ICP. | The value for zinc in the vendor stock solution has been determined to be 7% higher than indicated on the certificate of analysis. New calibration standards will be prepared using a new stock solution and verified by the end of December 2005. All data analyzed since the use of this stock have been within control. No corrective action is necessary on past data. | We will continue to monitor the zinc line for bias. If the new standards have not fixed the bias we will look for other causes. |
| Zinc, Whole-Water Recoverable ICP-MS labcode 2390 parmcode 1092D | This method has been discontinued by the NWQL. The replacement method is zinc, whole-water recoverable, cICP-MS (collision cell), labcode 3139, parmcode 1092E. |  |  |
| Previous comments for zinc, whole-water recoverable, ICP-MS, labcode 2390, parmcode 1092D |
| From QADATA Summary December 2005 | Data still exhibit a negative bias. Results from samples analyzed August 2005 and later are biased on average 8.7% below the target value. This bias is not evident for zinc determinations by ICP. | The value for zinc in the vendor stock solution has been determined to be 7% higher than indicated on the certificate of analysis. New calibration standards will be prepared using a new stock solution and verified by the end of December 2005. All data analyzed since the use of this stock have been within control. No corrective action is necessary on past data. | We will continue to monitor the zinc line for bias. If the new standards have not fixed the bias we will look for other causes. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Filtered Electrometric labcode 2109 parmcode 29801A | The data show a slight upward trend starting around July 2005 and continuing through December. The most recent results do not exhibit this bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| Alkalinity, Whole-Water Recoverable Electrometric labcode 0070 parmcode 90410A | The data show a slight upward trend starting mid-October 2005 and continuing through mid-January 2006. The most recent results do not exhibit this bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |