##### QADATA Summary – June 2006

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 100 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 11/01/05 through 05/25/06; however, longer-range charts may be used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the laboratory are removed from the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and how the corrective actions have impacted the new data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/FY06charts.htm.  Laboratory and district personnel can also make their own custom QADATA retrieval to parallel their objectives by using the QADATA Application online at  http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Nutrients

Phosphorus, Filtered,  (Low-level), Colorimetry, Labcode 2331, Parmcode 0666G	4

Trace Metals

Strontium, Filtered,  ICP-MS, Labcode 2507, Parmcode 1080I	5

Majors

Alkalinity, Filtered, Electrometry, Labcode 2109, Parmcode 29801A	6

Alkalinity, Whole-Water Recoverable, Electrometry, Labcode 0070, Parmcode 90410A	6

Fluoride, Filtered, ISE, Labcode 0031, Parmcode 0950B	6

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Phosphorus, filtered, colorimetry, low-level, parmcode 0666G – Increased variability, but less bias, starting January 2006.

- Alkalinity, filtered, electrometry,  parmcode 29801A – Positive bias starting October 2005.  Recent data show 7.8% bias.

- Alkalinity, whole-water recoverable, electrometry,  parmcode 90410A – Positive bias starting September 2005.  Recent data show 7.5% bias.

- Fluoride, filtered, ISE, parmcode 0950B – Increase in positive bias starting April 2006.  Data since then show 18% bias

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Phosphorus, Filtered Low-level Colorimetry labcode 2331 parmcode 0666G | A positive bias starting in 2005 is shown improving in February 2006. However, variability has increased as of late-January 2006. Since then, four out of 34 results (12%) are more than 4 f-pseudosigmas lower than the target value. | One of the samples was determined to be a mispour by the analyst. Further investigation is now under way to determine the cause of the lower-than-target values for the other three samples. | The analyst operating the instrument has been instructed to take greater care when preparing samples for analysis. We are examining the original data and are rerunning the other three samples to determine if they were mispoured as well. The examination is targeted to be completed by June 16, 2006. Further corrective action may be necessary depending on the results of the investigation. |

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Strontium, Filtered ICP-MS labcode 2507 parmcode 1080I | Results from the low concentration blind sample (at 31.8 ug/L) are biased on average 6.4% lower than the target value. In addition, about one-half of the results for this sample are outside of the lower control limit (-2 f-pseudosigmas). Results for all other blind samples submitted show acceptable bias and variability. | New calibration standards were prepared in May. All blind samples have been rerun using the new calibration standards and no bias is evident at this time. | Analysis will be closely monitored until next QADATA Summary is distributed. |
| Previous comments for strontium, filtered, ICP-MS, labcode 2507, parmcode 1080I |
| From QADATA Summary April 2006 | The data are now biased less than 5% from their target value. However, three out of the last 13 results (23%) are outside -2 f-pseudosigmas. Analysis should be closely monitored until the next QADATA Summary is distributed. |  |  |
| From QADATA Summary February 2006 | An upward trend in the data is shown starting December 2005. Data since then exhibit a positive bias of 7.1% above the target value. | There was increased scatter starting in October and the upward trend starting in December. Starting in Oct. most of the Sr results were from the new ICP-MS. | We transferred the Sr analysis to the older ICP-MS the week of January 29th and have stopped running Sr on the new instrument. There is not enough data to see the effect on the bias. If that fixes the problem we will not analyze Sr on the new instrument until we determine the cause of the problem and fix it. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Filtered Electrometric labcode 2109 parmcode 29801A | The data indicate a positive bias. Results from samples analyzed as of November 2005 are biased on average 7.8% above the target value. | Use of new instrumentation was initiated in July 2005. Prior to the new instrumentation, spike recoveries showed a negative bias. After the new instrumentation was initiated the bias became positive with decreasing magnitude and more scatter. Recoveries for reference material M-152 (target value = 6.00 mg/L) from April 2006 to present show a 0.2% positive bias. M-168 recoveries (target value = 28.2 mg/L) from March to present show a 3.4% positive bias. | At this time no corrective action will be taken. The inflection point technique for determining concentration is being investigated for use in this method. The investigation is targeted to be completed by October 1, 2006. |
| Alkalinity, Whole-Water Recoverable Electrometric labcode 0070 parmcode 90410A | The data indicate a positive bias. Results from samples analyzed as of November 2005 are biased on average 7.5% above the target value. | Use of new instrumentation was initiated in July 2005. Prior to the new instrumentation, spike recoveries showed a negative bias. After the new instrumentation was initiated the bias became positive with decreasing magnitude and more scatter. Recoveries for reference material M-152 (target value = 6.00 mg/L) from April 2006 to present show a 0.2% positive bias. M-168 recoveries (target value = 28.2 mg/L) from March to present show a 3.4% positive bias. | At this time no corrective action will be taken. The inflection point technique for determining concentration is being investigated for use in this method. The investigation is targeted to be completed by October 1, 2006. |
| Fluoride, Filtered ISE labcode 0031 parmcode 0950B | The data exhibit a marked upward shift starting April 2006 with some improve-ment shown in May. Results from samples analyzed as April are biased on average 18% above the target value. | The upward shift in bias seen in the BQS charts was not evident in the online QC charts for the months of April and May. The analytical system, however, has been thoroughly cleaned several times and all pump tubing replaced. A new reference and fluoride electrode were installed on May 17. | The high bias evident in April seems to have diminished. The Nutrient Unit will closely monitor line performance and BQS charts to see what effect the change of electrodes has made. If necessary, new standards will be made. |