Title: V9.1 SAS System Output
Author: SAS Version 9.1
Last saved by: dbeaty

10:00  Monday, November 13, 2006  

FY06 LT-MDL Review – Kone

10:00  Monday, November 13, 2006  

LTMDL Review for Ammonia, Kone

10:00  Monday, November 13, 2006

10:00  Monday, November 13, 2006  

LTMDL Review for Nitrite, Kone

10:00  Monday, November 13, 2006  

LTMDL Review for ortho-PO4, Kone

Analyte | TestID | Lab Code | Calculated Spike LT-MDL | Calculated Blank LT-MDL | Blank Corrected LT-MDL | LT-MDL in use | Update LT-MDL To: | Update LRL To: | Update?
Ammonia | 00608J | 3116 | 0.00703 | 0.01650 | 0.00793 | 0.005 | 0.0165 | 0.033 | Y
Nitrite | 00613I | 3117 | 0.00094 | 0.00099 | 0.00055 | 0.001 | | | N
ortho-PO4 | 00671L | 3118 | 0.00684 | 0.00112 | 0.00289 | 0.003 | ~~0.0068~~ | ~~0.014~~ | N

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
00608J | 3116 | 066108W | 14 | 0.01 | 0 | 0.4538 | Normal
| | 06Bk | 11 | . | 0 | 0.7294 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
00613I | 3117 | 066108W | 14 | 0.002 | 0 | 0.4257 | Normal
| | 06Bk | 12 | . | 0 | 0.1224 | Normal

TestID | Labcode | Solution | N | Expected | False Neg | P-value | distribution
00671L | 3118 | 066108W | 13 | 0.006 | 0 | 0.4539 | Normal
| | 06Bk | 12 | . | 0 | 0.2517 | Normal