# STEREO/IMPACT/SEP Operations Manual

# for the

# High Energy Telescope (HET)

# and

# Suprathermal Ion Telescope (SIT)

# Version 1.0

# 10/11/06

## Contributions by:

Glenn Mason

Tom Nolan

Don Reames

Larry Ryan

Peter Walpole

Kristin Wortman

Tycho von Rosenvinge

## Compiled by:

Kristin Wortman

# 1.0 Background

The STEREO/IMPACT/ HET and SIT flight software was designed to run on a CPU24 processor.  The HET interfaces with its front-end electronics and the SEP Central MISC through a series of Universal Asynchronous Receiver/Transmitters (UARTs).

## 1.1 System Overview

The LET, HET and SIT sensors each require a dedicated microprocessor for onboard data processing.  The microprocessor used for LET shall be the P24 MISC (Minimal Instruction Set Computer. The microprocessor used for HET shall be the CPU24 (24-Bit Embedded Microprocessor), which is described below and in Reference 3.  Processed data from all of the microprocessors associated with these three sensors shall be gathered by the SEP Central MISC (P24 processor), and formatted for transmission to the IMPACT DPU (per Reference 5). The SEPT sensor does not have a dedicated microprocessor, and data from SEPT shall flow directly to the SEP Central MISC. Some processing of SEPT data shall occur in the SEP Central MISC before the data are formatted and transmitted to the IMPACT DPU. Figure 1 shows a block diagram of the SEP Instrument Suite.

Figure 1: SEP Instrument Suite Block Diagram

## 1.2 Boot ROM

The CPU24 contains 16 words of ROM that holds a small program to boot over the serial link.  After a reset, the CPU24 starts execution at location 0 which is mapped to the internal ROM.  The instruction at location 0 performs a jump to location 20001 (hex) where the remainder of the boot code is located.  A fixed number of bytes are received over the serial link. Every three bytes received are packed into a 24-bit word, with the first byte going into the most significant slot and the third byte going into the least significant slot. Words are stored beginning at address 1 in SRAM. Execution begins at address 1 (in SRAM) following the serial transmission.

Telemetry mode.  The telemetry mode is a single byte that determines the settings of all five telemetry control parameters.  A mode switch is performed only on major frame boundaries.  When a new mode is commanded, the telemetry control parameters are copied from a set of hard-coded values associated with the mode.  The hard-coded values associated with each mode are set forth in the following table.

| Mode | Minor Frame Interval | Pkts per Minor Frame | Pkts per Major Frame | Defined Packet Sequence - HET | Description |
| --- | --- | --- | --- | --- | --- |
| 0 | 3 | 1 | 8 | G, F, A, B, C, C, C, D | Flight telemetry |
| 3 | 3 | 1 | 8 | A, B, G, F and Raw Event (ApId 597d) | Diagnostic flight mode |

## 1.3 Document Referenced

- Description of the Suprathermal-Ion-Telescope, Glenn Mason APL

- HET Telemetry Formatting, Don Reames, Kristin Wortman

- STEREO/IMPACT/SEP HET and SIT Flight Software, Tom Nolan, Kristin Wortman

- SIT Comprehensive Performance Test, Peter Walpole

- HET Comprehensive Performance Test, Larry Ryan,, Tycho von Rosenvinge, Kristin Wortman

- SIT HVPS In-flight Turn-on Procedure, Glenn Mason, Peter Walpole, Kristin Wortman

# Instrument Commanding

Commands are received from SEP Central over the bi-directional serial port. The transmission scheme is asynchronous, character-by-character.  The baud rate is 57600.  After the transmission of a command string is complete, SEP Central listens.  This is the window in which HET/SIT must transmit a response, which is a printable ASCII string.   After the response, the flight instrument issues its prompt (either “HET> ” or “SIT> ”). Commands may be received at any time, but they are executed only on major frame boundaries.

The flight software maintains a circular queue of incoming characters.  The arrival of a character on the bi-directional serial port triggers the highest-priority interrupt in the flight processor.  The interrupt service routine simply reads the character from the serial port and writes it to the queue.  No interpretation is done at this stage.  The queue can hold 2048 characters.

The background routine regularly checks the incoming character queue.  When a character is present, it is dequeued and passed to a finite-state machine.  The task of the finite-state machine is to recognize complete commands, issue a response to SEP Central, and place completed commands in another queue for execution at the next major frame.

Note that the command keywords “HETCMD” and “HETBINARY” which precede any sequence of commands routed through SEP Central are interpreted by SEP Central and not passed through to the instrument.  (In order to maintain compatibility during instrument testing, these commands actually are recognized by HET and treated as a no-op, with no echo.)

## 2.1 ASCII Command Processing

Except for binary commands (described below), all commands are in printable ASCII. They may be terminated by either a carriage return (13) or a newline (10).  The command state machine recognizes a complete command when one of these two characters is received.  It then performs command preprocessing. The command preprocessor takes the following steps:

- If the command string is empty (only a carriage return or newline received), then the instrument prompt is sent.

- If the command is one that requires immediate execution, it is dispatched to the command handler.  Immediate execution commands are peekw, immed, and dump.

- If the command string starts with a valid command keyword, then it is queued for later execution, and a command echo is sent, followed by the instrument prompt.

- If none of the above is true, an error message is sent and prompt is sent.

For valid commands, the command echo consists of a 4-digit identifier number (in hex) followed by a copy of the command string.  The first two digits of the identifier number are the least-significant byte of the current major frame number, and the second two digits are the command sequence number that is reset to zero in each major frame.  After the identifier, the command string is echoed as it was received.  The “cmdstr” command has a special echo to avoid repeating the entire 108-character string.

In immediate execution mode (immed 1), commands are queued for processing, but they are pulled off the queue and executed immediately instead of waiting for a major frame.  The command echo contains an asterisk inserted between the identifier number and the command echo.  This asterisk serves as a reminder that the command is executed immediately instead of delayed until major frame time.

When a command does not begin with a recognized keyword, an error message is sent, consisting of the received command followed by a question mark.  No identification number is sent in this case.

The final portion of the command echo is the instrument prompt, which is sent in all cases (delayed command, immediate command, or error).

## 2.2 Binary Command Processing

A binary command consists of an ASCII introducer followed by a binary load package.  The introducer is the ASCII command keyword  “binary.”  This introducer is followed by a carriage return or line feed, just like an ASCII command.  Binary commands do not go through the preprocessing described above.  Instead, when the state machine recognizes the binary introducer keyword, it prepares to receive a binary load package.

The binary load package contains a two-byte length, which is the number of bytes to follow (including checksum), then the command bytes themselves, then a two-byte checksum. Two-byte numbers are transmitted most-significant byte first.  The checksum is the 16-bit sum of all of the 8-bit command bytes.  The following table depicts the binary load package format.

| 0 | Length msb |
| --- | --- |
| 1 | Length lsb |
| 2, . . ., N-1 | N-2 command bytes, where N = Length |
| N | Checksum msb |
| N+1 | Checksum lsb |

When the finite-state machine recognizes the “binary” command introducer, it does not dispatch it to the command preprocessor as it would for an ASCII command. Instead, it enters a series of states to receive the length, the command bytes, and the checksum.  The binary command load is placed in a large memory buffer (the binary command load staging area).

When the two-byte checksum is received, the binary command payload is finished.  An echo is sent in one of the following two forms:

binary A:aaaaaaaa N:nnnnnnnn OK

or

binary A:aaaaaaaa N:nnnnnnnn ckserr cccccccc dddddddd

If the computed checksum matches the received checksum, the OK echo is sent.  The number following the “A:” is the relative address (i.e., the byte number within the command load staging area where the binary command payload was placed), and the number following the “N:” is the number of bytes in the command payload.  If the checksums do not match, then the ckserr echo is sent, followed by the received checksum and then the computed checksum.

A complete binary command load ordinarily consists of one or more “binary” commands with binary payloads, followed by a single ASCII “load” command.  The “load” command instructs the flight instrument to copy the binary load from the command load staging area to a specific address in memory, which is the absolute address of the table or memory area that is the ultimate destination of the command load (its target address).  When large tables are uploaded, multiple “binary” commands are generally sent before the “load” is issued.  The data in each binary command payload is copied to consecutive locations in the command load staging area.  The command echo verifies that this is the case:  in each successive echo the address will be incremented by the length of the previous load.

When a command load checksum fails, the relative address for the next load package is still incremented as though the load package were received correctly.  This is done so that a large number of command loads may be transmitted, the resulting command echoes examined, and only those loads that were not received correctly may be retransmitted.  The retransmission of a failed command packet must be preceded by a “loadat” command.  The “loadat” command allows the specification of the relative address for the retransmitted command load (which can be extracted from the command echo corresponding to the failed command load).

The “load” command has three options.  Load type zero (24-bit words) packs three bytes into each word at the target address.  Bytes are packed msb, middle, lsb.  The length of the table that is actually loaded at the target address is n/3, where n is the number of binary command bytes sent during the preceding binary commands.  Load type 1 (1 byte per word) copies one byte per word to the target address.  The two most-significant bytes of each target word are set to zero.  Load type 2 (2 bytes per word) packs two bytes into each word at the target address.  Bytes are packed msb, lsb (the most-significant byte of each 24-bit word at the target address is set to zero).  The length of the table that is actually loaded at the target address is n/3, where n is the number of binary command bytes sent.  The load type 1 and 2 options afford some measure of compression, because data stored at three bytes per word can be expanded to either one or two bytes per word as appropriate.

For the “load” command, the number of bytes transferred to the target address is determined by the preceding binary command sequence.  All bytes sent and received during the immediately preceding binary command load sequence (which may consist of one or more individual command loads) are copied into the target load address using the specified load method.  If the number of bytes is insufficient to fill out the last target word (for example, if an odd number of bytes were received and the load method is two bytes per word), the last word in the target area will have zeroes in the extra byte positions.  The “loadn” command accepts an additional parameter – the number of bytes to transfer – and otherwise operates just like the “load” command.  (If the byte count on the “loadn” command exceeds the number of uploaded bytes, the remainder will be undefined.)

After a “load” or “loadn” command, the flight software resets the relative address within the binary command staging area to zero in preparation for the next command load.  When a “load” command is sent with a target address of zero, the command has the effect of forcing the relative address to zero, but no data are transferred.  It would be good practice to issue a “load 0” command before beginning any binary load sequence because the “binary” commands rely on the flight software’s internal pointer mechanism.

The following command sequence illustrates the concepts described above to load a 1024-word table, where each word is 3 bytes.

| Command | Response | Comments |
| --- | --- | --- |
| load 0 | 0100 load 0 | Reset load address, prepare for binary load |
| binary |  | ASCII command, followed by CR or NL |
| [length=1026, 1024 data bytes, checksum] | Binary A:000000 N:000400 OK | Binary payload, length includes 2 checksum bytes |
| binary |  | ASCII command |
| [length=1026, 1024 data bytes, checksum] | Binary A:000400 N:000400 OK | Second portion of binary payload |
| binary |  | ASCII command |
| [length=1026, 1024 data bytes, checksum] | Binary A:000800 N:000400 OK | Third portion of binary payload |
| load 7000 0 | 0101 load 7000 0 | Transfer binary command load to absolute address 7000 (hex), 3 bytes per word at major frame 2 |

The following table shows the addresses where the HET tables will be loaded in RAM:

| HET Table Address (hex) | Length (hex) | Description |
| --- | --- | --- |
| 18000 | 4000 | Software counters (128 x 128) |
| 1C000 | 1000 | Particle type (64 x 64) |
| 1D000 | 20 | Addrtab (16x2) |
| 1D020 | 10 | Offsetch (8 x 2) |
| 1D030 | 10 | C-Delta E-Log GnFctr (8 x 2) |
| 1D040 | 10 | gfctrnum (8 x 2) |
| 1D050 | 2 | C-Resid E-Log H2 GnFctr (2) |
| 1D052 | 200 | C-Delta E-Log 2 |
| 1D252 | 200 | C-Resid E-Log 2 |

The following table shows the addresses where the SIT tables will be loaded in RAM:

| SIT Table Address (hex) | Length (hex) | Description |
| --- | --- | --- |
| 7000 | 800 | SSDHI |
| 7800 | 800 | SSDLO |
| 8000 | 4000 | BOX_ARRAY |
| C000 | 200 | TOFTAB |

### 2.2.1 Table Upload File Format for HET and SIT

For convenience, consistency, portability, and a host of other good reasons, a file format was developed to permit the contents of HET table uploads to be specified, edited, and packaged into binary commands for transmission to the spacecraft or instrument.  The flight software does not read or interact directly with table upload files in this format.  Instead, routines for reading the files are included in the instrument ground support systems.  Using these routines, the ground support systems read the table upload files and create binary commands in the format described above, plus applicable headers.  After mission-level, spacecraft-level, and SEP Central processing, the binary commands reach the flight software where they are interpreted and processed as described above to load the table contents into memory.

A table upload file is a readable ASCII file consisting of four kinds of lines: introducers, addresses, table contents, and comments.  They are described below.

An introducer line consists of the word “HETBINARY” or “SITBINARY” on a line by itself, with no spaces or other characters before or after it.  The purpose of having two different introducers is so that tables cannot be uploaded to the wrong instrument inadvertently.

An address line follows immediately after an introducer.  It contains three numbers, separated by spaces.  The format for numbers obeys C conventions.  That is, each number is assumed to be decimal, unless it begins with “0x,” in which case it is assumed to be hexadecimal.  The first number is the address at which the table is to be loaded.  This address is an absolute address in the flight software address space. The addresses of the major tables in the SIT and HET flight software are found in the discussion above.  The second number is the number of table entries to follow.  This normally corresponds to the size of the table.  The third number is the load type, which can be 0, 1, or 2.  Load type 0 treats each table entry as a 24-bit word. All three bytes of each word are transmitted to the flight software, then recombined into words and stored at successive address locations.  Load type 1 treats each table entry as an 8-bit value.  Only a single byte is transmitted to the flight software, which stores successive bytes at successive address locations, filling the most-significant 16 bits of each word with zeros.  If the table entry in the file is larger than can fit in one byte, it is truncated to 8 bits. Load type 2 treats each table entry as a 16-bit value.  Two bytes are transmitted to the flight software, which combines the two bytes into a 16-bit word, storing successive words at successive address locations, filling the most-significant 8 bits of each word with zeros.  The reason for having load types 1 and 2 is to conserve transmission bandwidth and load time when table entries are naturally limited in extent.

A table contents line consists of one or more numbers separated by commas, spaces, or tabs.  There must be the same number of table entries following the address line as specified in the second number of the address line (i.e., the table size).  However, the table entries can be spread out over any number of lines.  For example, each table entry could appear on a separate line, or all the table entries could appear on the same line (but lines are limited to 512 characters in total length).  As before, number formats obey C conventions.  They must begin with a digit, a minus sign, or a “0x.” Any other character appearing where a number ought to be ends that line (effectively acting as an inline comment).

A comment line is any line whose first character (other than spaces, tabs, or commas) is not a digit or a minus sign.  Comment lines can appear anywhere in the table upload file except between an introducer and the following address line.  The comment line immediately preceding the introducer is special – it is a shorthand description of the table that can be used by the upload software to inform the user of the table that is about to be uploaded.  Other comment lines appearing before the introducer are ignored, and any comment lines appearing in the table contents area are skipped until a line beginning with a numeric value is found.

Multiple table uploads can appear in one file.  Each table upload begins with an introducer and address line, and is followed by sufficient table contents lines to fill up the table. Following the table contents (and any optional comments), another introducer may begin a second table upload, and so on.

An example follows.

This table upload file contains two uploads.

First is a sample table containing 13 entries, where each entry is no larger than 16 bits.

HETBINARY

0x1f000 13 2

0, 10, 20, 50, 100, 200, 500, 1000 first 8 entries

2000, 5000, 10000, 20000, 50000 next 5 entries

Second is a sample table containing 4 entries, 24 bits each

HETBINARY

0x1f020 4 0

-1

-1

0x55aa55

-1

## 2.3 HET Command Descriptions

2.3.1 HET Commands

This section lists the HET common commands, the instrument-specific commands, and the table load addresses.  All commands are case-sensitive.  All numeric arguments must be supplied in hexadecimal.  A missing argument or a non-numeric argument will be interpreted as the number zero.

| Command | Description |
| --- | --- |
| tmode N | Set telemetry mode. N=0…5. See description above under Telemetry System. |
| modw A N | Modify word at address A to value N. |
| peekw A | Examine word at address A. |
| binary | Receive binary command load package. Note that this command is generated automatically by SEP Central when a HET-BINARY package is received. |
| immed N | Immediate command execution mode on (1) or off (0). Default for flight: immed off (0). |
| gwrite R N | Write value N to G-bus register R. [THIS MAY DISAPPEAR] |
| load A T | Transfer binary command load to address A using load method T (0=W24, 1=B, 2=W16). The number of bytes transferred is equal to the maximum relative address written since the last load or loadn command. This command is always executed immediately, and not delayed until the next major frame. The “dload” version of this command is to be used for delayed execution. |
| Dload A T | Same as load, except execution of the command is delayed until the next major frame. The “dload” version should be used to replace a table while the instrument is operational, to avoid corrupted major frames processed partially with the new table. The “load” version should be used for initial table loads to avoid overrunning the table buffer. |
| Loadn N A T | Transfer N bytes from binary command load staging area to address A using load method T (0=W24, 1=B, 2=W16). |
| loadat A | Set relative address to A for next binary command load. A is a byte number starting from zero within the current table. |
| Cgate N | Turn clock gating on (1) or off (0). |

2.3.2 HET PHASIC Chip Command Descriptions

There are two PHASIC chips in each HET flight unit, PHASIC 0 and PHASIC 1.  Both H1 detectors reside on PHASIC 0 and H2-H6 resides on PHASIC 1.  The channel settings used for each PHASIC chip is specific to the flight model and are described in the table below.

PHASIC Channel Configuration on HET Flight Boards

|  | Flight Model 1 | Flight Model 2 |
| --- | --- | --- |
| Detector | PHASIC | Channel | PHASIC | Channel |
| H1i | 0 | 1 | 0 | 1 |
| H1o | 0 | 11 | 0 | 10 |
| H2 | 1 | 3 | 1 | 2 |
| H3 | 1 | 5 | 1 | 5 |
| H4 | 1 | 7 | 1 | 7 |
| H5 | 1 | 9 | 1 | 10 |
| H6 | 1 | 12 | 1 | 13 |

The following HET commands configure the PHASIC chip registers, the onboard stimulus pulser, and the DAC settings.

| HET Command | Description |
| --- | --- |
| cmdstr N S | Command PHASIC N to string S. N=0, 1. S is 108 hex bytes long, representing left-justified 846-bit bit string. |
| dump N | Return contents of PHASIC N. |
| testp P A B | Write 14-bit period, 9-bit DAC A and 9-bit DAC B values to test pulser registers. The 9th bit of each DAC value is written to the respective DAC range bit. This command stops the automatic background test pulser algorithm. To resume automatic test pulser control, use “testp auto” command. |
| testp auto | Begin automatic background test pulser control algorithm (default). This command is used after a “testp” command has been sent to set manual values. |
| scopesel N V | Write 9-bit value V to scope select bits, PHASIC N |
| presel N V | Write 5-bit value V to preamp output mux select bits, PHASIC N |
| tpfbsel N P V | Test enable, test select, feedback select, PHASIC N (0-1), PHA P(0-15), value V (9 bits) |
| inresel N P V | Input resistor, input dac, PHASIC N (0-1), PHA P(0-15), value V (16 bits) |
| hgthrsel N P V | High gain threshold, PHASIC N (0-1), PHA P(0-15), value V (10 bits) |
| lgthrsel N P V | Low gain threshold, PHASIC N (0-1), PHA P(0-15), value V (10 bits) |
| phacont N P V | PHA control, PHASIC N (0-1), PHA P(0-15), value V (4 bits) |
| lgorsel N X V | rndn-gor-X L enable, X=0,1,2, PHASIC N (0-1), value V (16 bits) |
| hgorsel N X V | rndn-gor-X H enable, X=0,1,2 PHASIC N (0-1), value V (16 bits) |
| setdac V | Set housekeeping dac to value V |
| simevent N | Inject N simulated events into processing system. Event values are taken from loadable table. If table ends with –1, events are repeated until count expires. Issue “simevent 0” to resume normal external event processing. |

The following tables depict the commonly used PHASIC commands and the default boot flight configurations for all channels.

| Command | PHA chip | PHA channel | Hex bytes | # of bits | Description | Default in HET initial command string |
| --- | --- | --- | --- | --- | --- | --- |
| phacont | 0-1 | 0-F | x | 4 | hg-adc-en, lg-adc-en, verbose, power-on | Channels in use set to d, all unused channels set to 0 |
| tpfbsel | 0-1 | 0-F | xxx | 9 | test-enable, testsel3,2,1,0, fbsel3,2,1,0 | H1i, H1o, H2, H6 set to 6; H3, H4, H5 set to C |
| inresel | 0-1 | 0-F | xxxx | 13 | fdgsel2,1,0, indac9,…,0 | H1i, Hio, H2, H6 set to 1c30; H3, H4, H5 set to 1c60 |
| hgthrsel | 0-1 | 0-F | xxx | 10 | ioff9,…,0 (for high gain channel) | All channels set to provide lowest possible threshold |
| lgthrsel | 0-1 | 0-F | xxx | 10 | ioff9,…,0 (for low gain channel) | All channels set to provide lowest possible threshold |
| scopesel | 0-1 |  | xxx | 9 | scopesel7,…,0, scope-enable | 000 |
| presel | 0-1 |  | xx | 5 | presel3,2,1,0, preout-enable | 00 |

‘phacont’ Command

The ‘phacont’ command is used like this:

phacont  x  y  z

x = PHASIC # (0 or 1)

y = channel # (0 to F)

z = see table below for most common values to use

| z | Description |
| --- | --- |
| 0 | Turn off channel completely |
| 1 | Disable HG and LG ADC’s, but leave channel turned on |
| 7 | Disable HG ADC, enable LG adc |
| b | Enable HG ADC, disable LG adc |
| d | Enable HG and LG ADCs, verbose mode off (default at boot-up) |
| f | Enable HG and LG ADCs, verbose mode on |

‘tpfbsel’ Command

The ‘tpfbsel’ command is used like this:

tpfbsel  x  y  z

x = PHASIC # (0 or 1)

y = channel # (0 to F)

z = see table below for most common values to use

| z | Description |
| --- | --- |
| 6 | Disable test pulser, set feedback capacitor to 30pF, and set test input capacitor to 0pF (default at boot-up for H1i, H1o, H2, H6) |
| c | Disable test pulser, set feedback capacitor to 60pF, and set test input capacitor to 0pF (default at boot-up for H3, H4, H5) |
| 166 | Enable test pulser, set feedback capacitor to 30pF, and set test input capacitor to 18pF |
| 1cc | Enable test pulser, set feedback capacitor to 60pF, and set test input capacitor to 36pF |

Note that currently the software boots up with the test pulsers disabled and the test input capacitors set to 0.  For flight, you will want the test pulsers enabled and the test input capacitors set to the same hex setting as the feedback capacitors (6 for H1i, H1o, H2, H6; c for H3, H4, H5).  Using the same hex setting for feedback and test input capacitors allows the test pulser DAC to probe the full dynamic range of the HG and LG channels.  More on that below in the description of the ‘testp’ command.

‘inresel’ Command

The ‘inresel’ command should generally not be used, since the software automatically sets the indac values (input current DAC to balance detector leakage current).

‘hgthrsel’ and ‘lgthrsel’ Commands

These commands will generally not be used, since the thresholds have been set.  The current threshold values are programmed into the software and will display on the housekeeping page.  They should only be changed (increased) if the noise goes up.

‘scopesel’ and ‘presel’ Commands

These commands are only useful on the Engineering Model now, and they are described in Rick Cook’s PHASIC document.

‘testp’ Command

The ‘testp’ command is used to run the on-board test pulser, and its format is as follows:

testp  x  y  z

x sets the test pulser frequency:  freq = 1 / ( x/12500 + 1e-5 )  pulses per second

(the default at boot-up is currently 4e2 = 10 pulses per second)

y sets the DAC setting for PHA #0 (common test input to H1i and H1o)

z sets the DAC setting for PHA #1 (common test intput to H2, H3, H4, H5, H6)

(the defaults at boot-up for y and z are 0)

The DAC settings (y and z) can be as follows:

00 to ff  (DAC voltages = 0V to .25V)

100 to 1ff  (DAC voltages = 0V to 5.00V)

If the test input capacitor and feedback capacitor are set to the same hex setting (as described above in the ‘tpfbsel’ command section), the low range of DAC voltages will stimulate the entire HG channel, and the high range of DAC voltages will stimulate the entire LG channel.

The test pulser can be run in automatic mode, using this command:  testp auto

This will step through each of the DAC settings (one step per minute) at the last selected pulser frequency.

## 2.4 Software Status Indicators

The flight software maintains a number of counters, status flags, error indicators, and parameters as part of its operation.  Certain of these values are placed in the telemetry in each major frame as a check on flight software operation.  The tables below list each of the software telemetry points for HET and SIT.

| HET Variable | Location | Description |
| --- | --- | --- |
| CoincidenceCnt | Pkt 590, bytes 20-21 | Counts every event interrupt except stop_ev (g@8 bit 0 is set) and no_event (g@9 bit 8 is set). Cleared every minute. |
| lastcmd | Pkt 590, byte 44 | Highest command sequence number reached in major frame (equal to number of commands received that pass parsing). Resets from zero every minute. |
| CmdErrBits | Pkt 591, bytes 46-47 | Bit set when command parsed correctly but did not execute. Bit number corresponds to sequence number associated with command when received during previous major frame. Whenever any bit is set in this word, there will also be a software error flag set in bit 6 of Errflags (see below). |
| IdleCnt | Pkt 591, bytes 48-49 | Counts every pass through the background loop, read out and cleared every minute. Provides a “livetime” check, since more time spent in interrupt and event processing will decrease this count. For this check to work, clock gating needs to be off (cgate 0). |
| Errflags | Pkt 598, bytes 29-30 | Bit set corresponds to an error encountered during previous major frame time. Cleared every minute. Bit 0: Receive queue full Bit 1: Transmit queue full Bit 2: Command queue full Bit 3: Command buffer overflow Bit 4: Command handler timeout Bit 5: Command syntax error (command rejected) Bit 6: Command processing error (CmdErrBit set) Bit 7: Callback timer error Bit 8: ADC timed out Bit 9: Queuing error, queue reset to empty state |
| swVer | Pkt 598, bytes 31-32 | Software version tag. LSB (byte 31) is day of month, MSB (byte 32) is month in which version was created. |
| Err_event | Pkt 598, bytes 33-34 | Event read from hardware but no token bit was set. Event was counted in CoincidenceCnt but not processed. |
| No_event | Pkt 598, bytes 35-36 | Event interrupt handled but no_event bit was set to 1 (g@9 bit 8). Event not counted in CoincidenceCnt. |
| Lost_event | Pkt 598, bytes 37-38 | Event could not fit in raw event queue. The raw event queue is only used in laboratory testing. Data in raw event queue is used to fill packet 598. If this counter still exists in the flight version, it will increment at high data rates because the raw event queue is never emptied. However, it is not an error and the events are also processed normally. |
| tmMajorFrame | Pkt 598, bytes 39-40 | Major frame number. This is necessary in order to correlate housekeeping packets with science packets. Each science packet has the major frame number in the HET header following the CCSDS header. The housekeeping packet is merged into the SEP Central housekeeping and does not have a HET header. |
| checksum_csum | Pkt 598, bytes 41-43 | Three-byte checksum of entire table area. The table checksum is computed on the ground during the flight software build process, and should match this computed checksum. The checksum is computed one word at a time in the background idle loop. This may take some time if clock gating is enabled (cgate 1). Table entries are added using a 24-bit sum, ignoring carries. Once the entire table area has been added up, the result is telemetered and the calculation is repeated. |

## 2.5 SIT Command Descriptions

The following commands are passed on to the front end logic:

### 2.5.1	One-Bit Commands

The following commands consist of a single bit each and control a single function in the front-end logic.

- hvenable  - enables the high voltage power supply

1=enable, 0 = disable, turn-on state = 0,

Expected use:  sent every time instrument is turned on to NORMAL (or Science) mode.

- eonly - allows analysis of events based on energy without the TOF

1 = EONLY, 0 = normal mode, turn-on state = 0

Expected use - diagnostic or in case of failure of TOF system

2.5.2	Data Commands (8 bits)

The following commands contain 8 bits of data each and control an analog function in the front-end electronics.

- hvlevel - sets the top voltage out of the HVPS

values: 0-255, 0 = 0 volts, 255 = tbd (~5000v), turn-on state = 0

Expected use - several commands will be sent each time the instrument is turned on into NORMAL (or Science) mode to step the HV up to the correct operating level.  In addition, on rare occasions (perhaps once per year) the HV will need to be changed to compensate for operational loss of gain in the micro-channel plates.

- threshold - sets the SSD discriminator threshold

values: 0-255, 0 = low threshold, 255 = high threshold, turn-on state = 0

Expected use - send once every time the instrument is turned on.  On rare occasions it may be necessary to change the threshold to compensate for increased noise in the solid-state detector.

2.5.3 Commands to the CPU24

The following commands are processed within the CPU24, setting flags or changing values in memory.

2.5.3.1 State Command

The following command sets the CPU24 operating mode

- toferror - tells CPU24 whether to process events with TOF error bits set

1 = process events independent of TOF error bits

0 = only process events with TOF error flags = 0

turn-on state = 0

Expected usage - diagnostic and error recovery

      - Binary Commands

Refer to section 2.2 for a description on the binary commands to change the contents of CPU24 memory locations which store the SIT event-processing look-up tables.

## SIT Command Summary

The following tables describe the format of the SIT commands.

| SIT Command | Description |
| --- | --- |
| tmode N | Set telemetry mode. N=0…5. See telemetry mode descriptions in section 1.2. |
| modw A N | Modify word at address A to value N. |
| peekw A | Examine word at address A. |
| binary | Receive binary command load package. Note that this command is generated automatically by SEP Central when a SIT-BINARY package is received. |
| immed N | Immediate command execution mode on (1) or off (0). Default for flight: immed off (0). |
| gwrite R N | Write value N to G-bus register R. [THIS MAY DISAPPEAR] |
| load A T | Transfer binary command load to address A using load method T (0=W24, 1=B, 2=W16). The number of bytes transferred is equal to the maximum relative address written since the last load or loadn command. This command is always executed immediately, and not delayed until the next major frame. The “dload” version of this command is to be used for delayed execution. |
| dload A T | Same as load, except execution of the command is delayed until the next major frame. The “dload” version should be used to replace a table while the instrument is operational, to avoid corrupted major frames processed partially with the new table. The “load” version should be used for initial table loads to avoid overrunning the table buffer. |
| loadn N A T | Transfer N bytes from binary command load staging area to address A using load method T (0=W24, 1=B, 2=W16). |
| loadat A | Set relative address to A for next binary command load. A is a byte number starting from zero within the current table. |
| cgate N | Turn clock gating on (1) or off (0). |

| SIT Command | Description |
| --- | --- |
| hvenable N | Enable high voltage (N=0,1) |
| eonly N | Control eonly bit (N=0,1) |
| hvlevel V | Set 8-bit HV level |
| toferror N | TOF error events (N=0,1) |
| limhi V | Set LIMHI to 8-bit value |
| junk N | Store junk events (N=0,1) |
| hvramp V | Ramps the HV up to the HV level |

## 2.7 Diagnostics

The software error flags are accumulated during each major frame, i.e., every occurrence of an error causes the corresponding bit to be or’ed in to the error flags word.  At the end of the major frame the bits are read out into the housekeeping telemetry and cleared.  The bit definitions for software error flags are described in the following table.

| Bit | Definition (bit = 1) |
| --- | --- |
| 0 | Receive character queue overflow |
| 1 | Transmit character queue overflow |
| 2 | Incoming command queue overflow |
| 3 | Buffer overflow in command preprocessing |
| 4 | Command handler timeout |
| 5 | Command syntax error |
| 6 | Command processing error |
| 7 | ADC callback timer error |
| 8 | ADC timeout in phasic baseline |

# 3.0 Telemetry Packet Formats

## 3.1 HET Telemetry Packet Formats

All rates in the SIT packets are log compressed from 24-bit to 16-bit quantities for telemetry according to the algorithm given in section 3.4.6.

The following table provides the Application Identifiers (ApId) that appear in the CCSDS header to identify HET science packets.

| APID | Description |
| --- | --- |
| 590 (Hex 24E) | Singles RATE |
| 591 (Hex 24F) | Status and Single PHA |
| 592 (Hex 250) | Stopping Particle PHA |
| 593 (Hex 251) | Penetrating Particle PHA |
| 594 (Hex 252) | Table Listing |
| 595-596 (Hex 253-254) | Not Used |
| 597 (Hex (255) | Tmode=3 Raw Events |
| 598 (Hex 256) | Housekeeping |
| 599 (Hex 257) | Beacon |

The telemetry output of the STEREO IMPACT HET telescope contains CCSDS packets in 7 different data formats:

- Rate packets

- Status and single PH packets

- Stopping particle PH packets

- Penetrating particle PH packets

- Table status dump

- Beacon packets

- Housekeeping data packets

After individual particle pulse height (PH) events are recorded by HET, the onboard processing algorithm identifies particle species and energies and bins the particles in “software rate counters,” as distinguished from hardware counters in the front-end electronics.  The identification of these counters is given in Appendix A.  In addition to binning all the particles, samples of the raw PH events are selected in 8 categories (see Appendix B) for inclusion in the telemetry stream. The format of the PH events themselves is given in B.2.  Note that PH events can vary in length from 2 to 16 bytes (always even).  All rates are log compressed from 24-bit to 16-bit quantities for telemetry according to the algorithm given in Appendix C.  Quantities longer than one byte are written into the packets least-significant byte first.

The following sections describe the formats of individual packet types.  In normal operation, HET generates 6 primary packets during a one-minute frame; these might be formatted as follows:  1 A, 1 B, 3 Cs and 1 Ds.  E packets are multiplexed out on at a rate that can be selected by command.  In flight, typically, E packets replace a PH packet once every 16 min to produce a complete dump every ~5 days.  F and G packets contribute the HET portions of the SEP beacon and housekeeping data.  Section 8 discusses an algorithm for selection of sample PHs to fill the PH packets.

The CCSDS header format is defined in the STEREO MOC to POC Interface Control Document.

### 3.1.2 ‘A’ or Rate packets (ApID: 590 (dec)  24e (hex)):

Rate packets are formatted as follows:

| Offset | Bytes | A-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 2 | Livetime |
| 18 | 2 | Trigger rate |
| 20 | 2 | Coincidence rate |
| 22 | 2 | Total number of events (excludes stimulus events) |
| 24 | 2 | N singles queued |
| 26 | 2 | N stopping events queued |
| 28 | 2 | N penetrating events queued |
| 30 | 2 | N stopping H |
| 32 | 2 | N stopping He |
| 34 | 2 | N stopping heavies |
| 36 | 2 | N penetrating H |
| 38 | 2 | N penetrating He |
| 40 | 2 | N penetrating heavies |
| 42 | 2 | N invalid events – out of sequence (reg and stim events ) |
| 44 | 2 | N invalid events – H1i & H1o, but not stimulus events |
| 46 | 2 | N invalid events – inconsistent dE/dx (reg and stim events) |
| 48 | 2 | N invalid events – H1 not first ph (reg and stim events) |
| 50 | 2 | N stimulus events (all types of stimulus events) |
| 52 | 12 | 6 background event bins 0-5 |
| 64 | 150 | 75 stopping event bins 6-80 |
| 214 | 16 | 8 penetrating event bins 81-88 |
| 230 | 26 | 13 single event bins 89-101 |
| 256 | 14 | 7 stimulus event bins 102-108 |
| 270 | 1 | TBD |
| 271 | 1 | checksum |

### 3.1.3 ‘B’ or Status and Single PH Packets (ApID: 591 (dec)  24f (hex)):

B packets are somewhat of a catchall.  They contain instrument status and health bytes, H1 single PH events, stimulator (STIM) event PHs and a few extra rates.  They are formatted as follows:

| Offset | Bytes | B-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 28 | 14 Single detector rates |
| 44 | 1 | number of commands received in previous major frame |
| 45 | 1 | zero |
| 46 | 2 | command errors (bit N = 1 if command N had an execution error, N=0-15) |
| 48 | 2 | background idle counts (compressed) |
| 50 | 14 | Offsets for current selected channels (h1i logain=0, h1i logain=1, h1o logain=0, h1o logain=1, h2 logain=0, h2 logain=1, h3 logain=0, h3 logain=1... h6 logain=0, h6 logain=1) |
| 64 | 7 | h1i chip=0 addr, h1o chip=0 addr, h2 chip=1, addr, h2 chip chip =1 addr..h6 chip=1 addr |
| 71 | 3 | Status Bytes (TBD) |
| 74 | 100 | 50 sample H1-only PHs |
| 174 | 96 | STIM PH events |
| 270 | 1 | N stimulus events in packet |
| 271 | 1 | Checksum |

### 3.1.4 ‘C’ or Stopping-Particle PH Packets (ApID: 592 (dec)  250 (hex)):

C packets contain PH events for stopping particles.  They are formatted as follows:

| Offset | Bytes | C-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 2 | N of stopping events |
| 18 | 252 | Stopping event PHs |
| 270 | 1 | TBD |
| 271 | 1 | checksum |

Stopping particles produce PHs in from 2 to 5 detectors, so their description contains from 6 to 12 bytes (appendix B).  This means that a packet can contains a maximum of 42 events, but contains at least 21 events if they are available.  Since event lengths are variable, extra space may exist at the end of the PH region that is too small for another event.  Any such bytes following the last PH event must be 0 filled.

### 3.1.5  ‘D’ or Penetrating-Particle PH Packets (ApID: 593 (dec)  251 (hex)):

D packets contain PH events for penetrating particles.  They are formatted as follows:

| Offset | Bytes | D-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 2 | N of penetrating events |
| 18 | 252 | 0-18 Penetrating event PHs |
| 270 | 1 | TBD |
| 271 | 1 | Checksum |

## Penetrating particles produce PHs in 6 detectors, so their description contains 14 bytes (appendix B).

### 3.1.6 ‘E’ or Table Listing Packets (ApID: 594 (dec)  252 (hex)):

E packets contain a listing of a segment of the table and constant region of the MISC-24 processor’s memory.  They are formatted as follows:

| Offset | Bytes | E-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 3 | Beginning address |
| 19 | 252 | Data |
| 271 | 1 | Checksum |

Generally, “Data” will consist of the next 84 24-bit words of memory beyond the beginning address.  These packets are designed to slowly multiplex the contents of large sections of memory into the telemetry stream.  Typically, one E-packet will be written every 16 major frames (minutes) in place of a PH packet.

### 3.1.7 ‘F’ or Beacon Packets (ApID: 599 (dec)  257 (hex)

Beacon packets transmit HETs share of the SEP beacon packet.

| Offset | Bytes | F-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 2 | 2 | Electrons 0.7-4 MeV - sum of sw bins 6-8 |
| 4 | 2 | Protons 13-21 MeV – sum of sw bins 9-12 |
| 6 | 2 | Protons 21-40 MeV – sum of sw bins 13-18 |
| 8 | 2 | Protons 40-100 MeV – sum of sw bins 81-82 |
| 10 | 2 | He 13-21 MeV/n – sum of sw bins 24-27 |
| 12 | 2 | He 21-40 MeV/n – sum of sw bins 20-22 |
| 14 | 2 | He 40-100 MeV/n – sum of sw bins 86-87 |
|  | 2 | C+O 30-52 MeV/n - sum of sw bins 35-39, 42-46 |
|  | 2 | C+O 52-74 MeV/n - sum of sw bins 40-41, 47-48 |
|  | 2 | Fe 52-74 MeV/n - sum of sw bins 73-74 |
|  | 2 | Livetime |
|  | 2 | Stop. efficiency (TBD) |
|  | 2 | Pen. efficiency (TBD) |
|  | 2 | HET status (TBD) |
| 270 | 1 | TBD |
| 271 | 1 | Checksum |

The rate quantities in the beacon data packet are derived by summing software rates described in Appendix A.   Regions in the packet other than the header and defined HET data block are filled with 0.

### 3.1.8 ‘G’ or Housekeeping Packets (ApID: 598 (dec)  256 (hex))

G packets contain the HET contribution to the housekeeping data packet.  Regions other than those defined are 0 filled.

| Offset | Bytes | G-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | ADC Temp 1 |
| 12 | 1 | ADC Temp 2 |
| 13 | 1 | PHASIC 0 PH channel ID |
| 14 | 1 | PHASIC 0 ADC Preamp |
| 15 | 2 | PHASIC 0 high gain threshold |
| 17 | 2 | PHASIC 0 low gain threshold |
| 19 | 2 | PHASIC 0 leakage current DAC setting |
| 21 | 1 | PHASIC 1 PH channel ID |
| 22 | 1 | PHASIC 1 ADC Preamp |
| 23 | 2 | PHASIC 1 high gain threshold |
| 25 | 2 | PHASIC 1 low gain threshold |
| 27 | 2 | PHASIC 1 leakage current DAC setting |
| 29 | 2 | Error Flags (16 bits) |
| 31 | 2 | Software version ID (16 bits) |
| 23 | 2 | N invalid token |
| 35 | 2 | N invalid trigger |
| 37 | 2 | N lost raw events |
| 39 | 2 | Major frame number |
| 41 | 3 | Table checksum |
| 44 | 1 | 24-bit DAC value, bits 0:7 (PHASIC 0 DAC) |
| 45 | 1 | 24-bit DAC value, bits 8:15 (PHASIC 1 DAC) |
| 46 | 1 | 24-bit DAC value, bits 16:23 (un:4, mux:2, rng1:1, rng0:1) |
| 47 | 5 | Available (TBD) |
| 52 | 219 | Not available for use |
| 271 | 1 | Checksum |

### 3.1.9 Raw Event Packet  (ApId: 597 (dec)  255 (hex))

This packet is a diagnostics packet and should not normally be generated in flight mode.

The raw event packet is generated when the FSW has been configured to tmode 3 when connected to SEP Central, and tmode 1 when not connected to SEP Central.  This packet replaces the packets with ApId 592 and 593 in tmode 3, and is generated in addition to the other packets at every unused second in tmode 1.  This packet should generally be used when analysis requires reviewing the raw (24 bit) events before particle processing.

| Offset | Bytes | 597-packet contents |
| --- | --- | --- |
| 0 | 11 | CCSDS header |
| 11 | 1 | HET mode byte |
| 12 | 2 | TBD |
| 14 | 2 | Major frame number |
| 16 | 255 | 85 24-bit raw events |
| 271 | 1 | Checksum |

## 3.2 HET Pulse Height Event Selection

In large solar energetic-particle events, the number of particles collected by HET will exceed by far the space available for PH events in the PH packets.  Since most of these particles are protons, intelligent sampling is required to insure that other species are sampled.  As a part of the onboard processing and binning of each particle, sampled PH events for stopping and penetrating particles are queued as protons, helium, or heavy ions.  Approximately 1/3 of the telemetry space will be reserved for each species, but vacant space will be filled.  That is, if there are 3 C packets being sent per frame, one will be dedicated to each species.  However, if the “heavies” packet is not full, the space will be filled by any left over He.  Then, if the Heavy and He packets have space, it will be filled with H PH events.  A similar process will be used to fill the 18 slots in the D packet with penetrating-ion events (6 each).

If all stopping PH events fit in 3 C packets, all will be sent; if all fit in 1 or 2 C packets only, those will be sent, leaving room for extra D packets.  In quiet times, 1 D packet should contain all Galactic Cosmic Ray events, but additional D packets might be useful very early in a large SEP event, before many stopping ions arrive.  When all pulse-heights fit in a 1 C packet and 1 D packet, 2 E packets may be sent to fill the 6 packet/min frame.  This is the minimum quiet-time set of packets.

Note that the queues for the 8 categories of sample PH events (see B.1) will have to be long enough to accommodate the maximum “backfill” of the packets.  For example the stopping proton queue must fill 3 C packets (max. 126 events) in case there are no stopping He & heavies; He must fill 2 C packets (max. 84 events), and heavies 1 (max 42 events).  Similarly, penetrating H events must fill as many as 3 D packets (54 events) and pen. He must fill 2 D packet (36 events) and heavies 1 (18 events).  Note that only those PHs to be actually transmitted need to be converted to compressed PH format.

Frame rates higher than one per minute may be required for accelerator calibrations.

## 3.3 HET Onboard Software Count Bins

### 3.3.1  H1 Singles Events

H1 – only events are mapped to MeV and binned in 2-MeV intervals out just beyond the proton endpoint, and then 4-MeV intervals to beyond the He endpoint.  Intervals (18) are:

0-2, 2-4, 4-6, 6-8, 8-10, 10-12, 12-14, 14-16, 16-20, 20-24, 24-28, 28-32, 32-36, 36-40, 40-44, 44-48, 48-52, >52.

### 3.3.2   Stopping Particles

Identified species and energy intervals are shown:

| MeV/n | H&4He | 3He | C | O | Ne | Mg | Si | Fe |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 13-15 | X |  |  |  |  |  |  |  |
| 15- | X |  |  |  |  |  |  |  |
| 17- | X | 17-21 |  |  |  |  |  |  |
| 19- | X |  |  |  |  |  |  |  |
| 21- | X | 21-27 |  |  |  |  |  |  |
| 24- | X |  |  |  |  |  |  |  |
| 27- | X | 27-33 | x |  |  |  |  |  |
| 30- | X |  | x | x |  |  |  |  |
| 33- | X | 33-40 | x | x | x |  |  |  |
| 36- | X |  | x | x | x |  |  |  |
| 40- |  | 40-47 | x | x | x | x | x |  |
| 45- |  |  | x | x | x | x | x |  |
| 52- |  |  | x | x | x | x | x | x |
| 62- |  |  | x | x | x | x | x | x |
| 74- |  |  |  | x | x | x | x | x |
| 87- |  |  |  |  | x | x | x | x |
| 98- |  |  |  |  |  | x | x | x |
| 109- |  |  |  |  |  |  | x | x |
| 119- |  |  |  |  |  |  |  | x |
| 140-163 |  |  |  |  |  |  |  | x |
| Bin-Count | 20 | 5 | 8 | 8 | 8 | 7 | 8 | 8 |

Total 72 bins for stopping ions, plus 3 background bins.

Electrons are accumulated in 3 bins:  0.7-1.4, 1.4-2.8, and 2.8-4.0 MeV

### 3.3.3  Penetrating Particles

Bins for penetrating particles are as follows:

| MeV/nucleon | H | He |
| --- | --- | --- |
| 40-60 | x | x |
| 60-100 | x | x |
| 100-200 | x | x |
| 200-400 | x |  |
| >400 | x |  |
| Bin-count | 5 | 3 |

There are a total of 8 bins for penetrating particles, plus 2 background bins.

### 3.3.4 Pulse-Height Selection and Formatting

### 3.3.4.1  Pulse-Height Categories

In the process of the onboard identification and binning of particle pulse-height (PH) events, sample events are selected and formatted for telemetry.   To prevent PH events of a given type (e.g. protons) from dominating the telemetry stream, events are grouped in the following categories:

0 - H1 singles

1 - Stopping protons

2 - Stopping He

3 - Stopping heavies

4 - Penetrating protons

5 - Penetrating He

6 - Penetrating heavies

7 - PH stimulator events

Singles events consist of a single PH in the H1 (H1i or H1i) detector.  Stopping events consist of PHs from 2 to 5 detectors, H1, H2, (…H5).  Penetrating events consist of 6 detectors H1 through H6.  PH stimulator events consist of up to 7 PHs: H1i, H1o, H2, …H6.  All events are presumed to be valid events, in that the PHs are ordered consistently, since invalid events (e.g. H1H2H5) have been weeded out early during PH interrupt service.  Each event has been tallied in an onboard particle species and energy “software” bin, although that bin may be a “background” bin between actual particle tracks.

Each event category is allocated a fixed basic amount of space in the output data packets.  This insures adequate representation of the categories in a large solar particle event.  However, if there are not enough events in a given category to fill the allotted space at the end of each one-minute “frame,” events from another category are allowed to occupy that space, in a priority order, until all available PH telemetry space in the packets is filled or all events are telemetered.  This means that more PH events in each category should be queued than can fit in the basic allocation for that category.  Basic allocations and filling priorities are presently TBD.

### 3.3.4.2  PH Event Format

With the exception of the H1 singles events, discussed below, all PH events consist of a 16-bit header followed by the appropriate number of 16-bit packed pulse heights.  The bit pattern of the 16-bit header is as follows (listed in lsb to msb order):

3-bits   Count of PHs in this event

8-bits   Onboard SW bin this event was assigned to

1-bit    Stimulator event flag

1-bit    Current rate mode of the HET

3-bits  PH category

Each individual PH is compressed from the 24-bit value read from the ASIC to a 16-bit value with the following bit pattern (lsb to msb)

11-bits  PH value

1-bit    Overflow bit

1-bit    High/low gain

3-bit     PH number  (H1i, H1o, H2, …H6)

To save space, H1 singles are treated differently from other events in that they have no header but consist only of a single 16-bit PH event as defined above.  Because of this, they occupy a fixed space in the telemetry packets that cannot be shared with events in other PH categories.

Note that stopping and penetrating events can be mixed arbitrarily.  Using the PH-counts in the PH-header it is possible to traverse logically from event to event throughout a list of events.

In the process of storing PH events in a 272-byte packet buffer, one may arrive at a place where an event is too large to fit in the remaining space.  In this case, the remaining space in the packet is zeroed and the event is stored elsewhere or omitted.  This will result in a PH-count of 0 when an attempt is made to read this event on the ground.

## 3.4  SIT Telemetry Packet Format Descriptions

All rates in the SIT packets are log compressed from 24-bit to 16-bit quantities for telemetry according to the algorithm given in section 3.4.6.

The following table provides the Application Identifiers (ApId) that appear in the CCSDS header to identify HET science packets.

| SIT APID | Description |
| --- | --- |
| 605 (Hex 25D) | RATE |
| 606 (Hex 25E) | PHA packet #1 |
| 607 (Hex 25F) | PHA packet #2 |
| 608 (Hex 260) | PHA packet #3 |
| 609 (Hex 261) | PHA packet #4 |
| 610 (Hex 262) | PHA packet #5 |
| 611 (Hex 263) | PHA packet #6 |
| 612 (Hex 264) | PHA packet #7 |
| 613 (Hex 265) | PHA packet #8 |
| 614 (Hex 266) | PHA packet #9 |
| 615 (Hex 267) | PHA packet #10 |
| 616 (Hex 268) | PHA packet #11 |
| 617 (Hex 269) | Test mode=2 (raw events) |
| 618 (Hex 26A) | Housekeeping |
| 619 (Hex 26B) | Beacon Rates |
| 623 (Hex 26F) | Fill |

### 3.4.1 Rate Packet (ApID: 605)

The rate packet contains discriminator and matrix rates, and command status information. There is no multiplexing.

| Byte # | Description |
| --- | --- |
| 1-11 | CCSDS |
| 12-13 | Discriminator Rate (=DR) 1-- START singles |
| 14-15 | DR2 – STOP singles |
| 16-17 | DR3 – Valid Stop |
| 18-19 | DR4 – SSD singles |
| 20-21 | DR5 – Event (triple coincidence) |
| 22-23 | DR6 – Dead time counter |
| 24-25 | DR7 – Artificial STOP count (TOF diagnostic) |
| 26-27 | DR8 – TOF system error count |
| 28-29 | Matrix Rate (=MR) MR1 |
| 30-31 | MR2 |
| 32-33 | MR3 |
| 34-35 | MR4 |
| 36-37 | MR5 |
| 38-259 | MR6 – MR116 |
| 260 | hvstep |
| 261 | 4 1-bit flags: bit 0 (lsb): 0 = TOF error events transmitted 0 = TOF error events dropped bit 1: 0 = HV disabled 1 = HV enabled bit 2: 0 = VS required for analysis 1 = SSD only required for analysis bit 3: 0 = ROM box 0 events dropped 1 = ROM box 0 events transmitted |
| 262-263 | LIMHI |
| 264-266 | 3-byte lookup table checksum |
| 267-271 | spare |
| 272 | checksum |

### PHA  Packets (ApIDs: 605-616)

| Byte # | Description |
| --- | --- |
| 1-11 | CCSDS Header |
| 12-15 | PHA event 1 |
| 16-19 | PHA event 2 |
| 20-23 | PHA event 3 |
| 24-27 | PHA event 4 |
| 28-267 | PHA event 5-64 |
| 268-270 | spare |
| 271 | Number of PHA events in packet |
| 272 | checksum |

### PHA  Raw Event Packet (ApID: 617)

| Byte # | Description |
| --- | --- |
| 1-11 | CCSDS Header |
| 12-15 | PHA event 1 |
| 16-271 | PHA event 2 - 65 |
| 272 | checksum |

### 3.4.4 Housekeeping Packet (ApId 618)

| Byte # | Description |
| --- | --- |
| 1-11 | CCSDS Header |
| 12-13 | Major Frame # |
| 14-15 | TOF gain Cal * 2048 |
| 16-17 | TOF Cal offset * -64 |
| 18 | TOF Cal error |
| 19 | HV monitor |
| 20 | TOF temp |
| 21 | SSD temp |
| 22 | foil temp |
| 23 | +3.3 V monitor |
| 24 | +2.4 V monitor |
| 25 | +5.0 Digital V monitor |
| 26 | +6.0 V monitor |
| 27-28 | Software version |
| 29-31 | lookup table checksum |
| 32-271 | unused |
| 272 | checksum |

### 3.4.5 Beacon Data Packet (ApId 619)

| Byte # | Description |
| --- | --- |
| 1-11 | CCSDS Header |
| 12-13 | Beacon Rate 1 (compressed) |
| 14-34 | Beacon Rate 2 – 12 (compressed) |
| 35-271 | unused |
| 272 | checksum |

### 3.4.6 .  Rate Compression Algorithm

/* 32-bit (or 24-bit) -> 16-bit compression for SW and HW rates  */

/* useage:  rateout=pack_rate(ratein);                           */

unsigned int pack_rate(ratein)

long ratein;

{

unsigned int rateout, power=0;

while (ratein&0xfffff000)

{

power+=0x0800;

ratein>>=1;

}

rateout=ratein;

if (power)

rateout=power+0x0800|((rateout&0x07ff));

return rateout;

}

/* Unpacking (not required in flight code) */

long long_rate(packed)  /* Unpack to long */

unsigned packed;

{

int power;

long out;

power= packed>>11;

if (power>1)

{

out=((packed&0x07ff)|0x0800);

out=out<<(power-1);

}

else

out=packed;

return out;

}

double dbl_rate(packed) /* Unpack to double */

unsigned packed;

{

int power;

double out;

power= packed>>11;

if (power>1)

{

out=((packed&0x07ff)|0x0800);

out=out*pow(2.,(double)(power-1));

}

else

out=packed;

return out;

}

# 4.0  Commanding and Test Procedures

## 4.1 HET Aliveness Procedure

To verify that HET has been booted up correctly, run the following aliveness procedure and view the packets on the GSE display.

### Record HK data for 7 detectors from HET GSE. It takes 5-6 minutes for data to come through after HET boots.

### Start Time: _____   HET Temperature #1 _____   HET Temperature #2 _____

Data from HET HK (dec) page on HET GSE:

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC |  |  |  |  |  |  |  |
| Channel |  |  |  |  |  |  |  |
| Preamp |  |  |  |  |  |  |  |
| HG Thresh |  |  |  |  |  |  |  |
| LG Thresh |  |  |  |  |  |  |  |
| Leakage |  |  |  |  |  |  |  |
| DAC |  |  |  |  |  |  |  |

The following are the expected HET HK values for HET FM1 and FM2. Compare them to the actual values above and pay attention to FM1 vs. FM2 values. Report discrepancies in the area below.

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC | 0 | 0 | 1 | 1 | 1 | 1 | 1 |
| Channel (FM1) | 1 | 11 | 3 | 5 | 7 | 9 | 12 |
| Channel (FM2) | 1 | 10 | 2 | 5 | 7 | 10 | 13 |
| Preamp (volts) | ~4.00 | ~4.00 | ~4.00 | ~4.00 | ~4.00 | ~4.00 | ~4.00 |
| HG Thresh (FM1) | 284 | 220 | 36 | 188 | 140 | 180 | 100 |
| HG Thresh (FM2) | 192 | 244 | 124 | 120 | 68 | 68 | 30 |
| LG Thresh (FM1) | 176 | 192 | 112 | 8 | 96 | 112 | 248 |
| LG Thresh (FM2) | 192 | 192 | 44 | 56 | 64 | 52 | 52 |
| Leakage (uA) | <0.5 | <0.5 | <0.5 | <2.0 | <2.0 | <2.0 | <1.0 |

______ UTC	Record the H1-H6 singles rates (low gain and high gain) for three consecutive major frames.  The high-gain singles rates should be < 100/min., and the low-gain singles rates should be < 5/min.

|  | High-Gain Singles Rates for 3 major frames | Low-Gain Singles Rates for 3 major frames |
| --- | --- | --- |
| Mjr Fr # |  |  |  |  |  |  |
| H1i |  |  |  |  |  |  |
| H1o |  |  |  |  |  |  |
| H2 |  |  |  |  |  |  |
| H3 |  |  |  |  |  |  |
| H4 |  |  |  |  |  |  |
| H5 |  |  |  |  |  |  |
| H6 |  |  |  |  |  |  |

Record the following rates for one major frame.    Major Frame #: ____________

| Rate Counter | Expected Counts | Actual Counts | Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- | --- | --- | --- |
| Livetime | ~12 x 106 |  | Stop heavies | 0 |  |
| Trigger | <300 |  | Pen H | 0 |  |
| Coincidence | 0 |  | Pen He | 0 |  |
| Events | 0 |  | Pen heavies | 0 |  |
| Singles que | 0 |  | Invalid | 0 |  |
| Stop que | 0 |  | H1 I and O | 0 |  |
| Pen que | 0 |  | Inconsist. | 0 |  |
| Stop H | 0 |  | H1 not first | 0 |  |
| Stop He | 0 |  | Stim evt cnt | 0 |  |

## 4.2 SIT Aliveness Procedure

To verify that SIT has been booted up correctly, run the following aliveness procedure and view the packets on the GSE display.

Verify SIT table checksums available in housekeeping data on SIT GSE:

| Description | Expected Value | Actual Value |
| --- | --- | --- |
| SIT FM1 tab cksm* | 52A82E (4/28/06) |  |
| SIT FM2 tab cksm* | 46E9B7 (4/28/06) |  |

*Note:  The table checksums will change when new lookup tables have been loaded.   Use the current table checksums to verify.  The checksums in the above table were correct  at the time of the STEREO launch.

Verify the following singles rates on SIT GSE match expected values:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | <20 |  |  |
| VSE | 0 |  |  |
| Matrix Rates (ApId 605) | All 0 |  |  |
| Beacon Rates (ApId 619) | All 0 |  |  |
| Pulse Height Events (ApId 606-617) | All 0 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected values:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SW Ver | 0903 |  |  |
| SW error | 0 = (no error) |  |  |
| HV Step | 00 |  |  |
| Toferror | 1 = error events |  |  |
| HV Status | 0 = off |  |  |
| Eonly | 0 = ET coincidence required |  |  |
| Junk (Events) | 0 = junk events ignored |  |  |
| LIMHI | 500 |  |  |
| calib gain | 9 - 11 |  |  |
| calib offset | -14 to –65, should settle at –15 but may take as much as an hour (don’t wait for it!). |  |  |
| calib error | 0x08, should settle to 0x00 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected values:

| Description | Expected Values | Actual Values | Comments |
| --- | --- | --- | --- |
| HV | -79.6 |  |  |
| Tof T | 25-35 C |  |  |
| Foil T | 20-25 C |  |  |
| SSD T | 20-25 C |  |  |
| +3.3 V monitor | 3.3 V |  |  |
| +2.5 V monitor | 2.5 V |  |  |
| +5 V monitor | 5 V |  |  |
| +6 V monitor | 6 V |  |  |

## 4.3   HET PHASIC Test Procedure Using Onboard Stimulus Pulser

### 4.3.1 Test Description

This comprehensive test procedure has been designed to run on either HET flight model (FM1 and FM2).  The WinMac-GSE software is required to send the appropriate commands for configuration and to view the HET packets to extract the relevant information for verification.

HET should be powered with the high voltage bias applied to the detectors, and the HET flight code should be booted with the HET tables successfully loaded.  The HET flight model should be in default boot configuration before running any test.  If previous tests were conducted, then HET should be rebooted before running any further tests.

The results of these tests will verify that the HET flight model is operational and the flight software is accepting and processing commands.

### Housekeeping Data

Record the housekeeping data for the seven HET detectors.  If HET has just been booted, it will take 5-7 minutes for all housekeeping data to come through.

Start Time: _____   HET Temperature #1 _____   HET Temperature #2 _____

Data from HK (dec) page:

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC |  |  |  |  |  |  |  |
| Channel |  |  |  |  |  |  |  |
| Preamp |  |  |  |  |  |  |  |
| HG Thresh |  |  |  |  |  |  |  |
| LG Thresh |  |  |  |  |  |  |  |
| Leakage |  |  |  |  |  |  |  |
| DAC |  |  |  |  |  |  |  |

The following are the expected values for FM1 and FM2:

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC | 0 | 0 | 1 | 1 | 1 | 1 | 1 |
| Channel (FM1) | 1 | 11 | 3 | 5 | 7 | 9 | 12 |
| Channel (FM2) | 1 | 10 | 2 | 5 | 7 | 10 | 13 |
| Preamp (volts) | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 |
| HG Thresh (FM1) | 284 | 220 | 36 | 188 | 140 | 180 | 100 |
| HG Thresh (FM2) | 192 | 244 | 124 | 120 | 68 | 68 | 30 |
| LG Thresh (FM1) | 176 | 192 | 112 | 8 | 96 | 112 | 248 |
| LG Thresh (FM2) | 192 | 192 | 44 | 56 | 64 | 52 | 52 |
| Leakage (uA) | <0.5 | <0.5 | <0.5 | <2.0 | <2.0 | <2.0 | <1.0 |

### Singles Rates / Spurious Events

Record the H1-H6 singles rates (low gain and high gain) for three consecutive major frames.  The high gain singles rates should be < 100/min., and the low gain singles rates should be < 5/min.

|  | High-Gain Singles Rates for 3 major frames | Low-Gain Singles Rates for 3 major frames |
| --- | --- | --- |
| Mjr Fr # |  |  |  |  |  |  |
| H1i |  |  |  |  |  |  |
| H1o |  |  |  |  |  |  |
| H2 |  |  |  |  |  |  |
| H3 |  |  |  |  |  |  |
| H4 |  |  |  |  |  |  |
| H5 |  |  |  |  |  |  |
| H6 |  |  |  |  |  |  |

Record the following rates for one major frame.    Major Frame #: ____________

| Rate Counter | Expected Counts | Actual Counts | Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- | --- | --- | --- |
| Livetime | ~12 x 106 |  | Stop heavies | 0 |  |
| Trigger | <300 |  | Pen H | 0 |  |
| Coincidence | 0 |  | Pen He | 0 |  |
| Events | 0 |  | Pen heavies | 0 |  |
| Singles que | 0 |  | Invalid | 0 |  |
| Stop que | 0 |  | H1 I and O | 0 |  |
| Pen que | 0 |  | Inconsist. | 0 |  |
| Stop H | 0 |  | H1 not first | 0 |  |
| Stop He | 0 |  | Stim evt cnt | 0 |  |

### Event Identification / Commanding

### Stopping H Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 1 6				tpfbsel 0 1 6

tpfbsel 1 5 c				tpfbsel 1 5 c

tpfbsel 1 7 c				tpfbsel 1 7 c

tpfbsel 1 9 c				tpfbsel 1 a c

tpfbsel 1 c 6				tpfbsel 1 d 6

tpfbsel 0 b 116			tpfbsel 0 a 116

tpfbsel 1 3 116			tpfbsel 1 2 116

testp 4e2 2d 1e			testp 4e2 2d 1e

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop H | ~600 |  |
| Stim evt cnt | ~600 |  |

### Stopping He Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 b 6				tpfbsel 0 a 6

tpfbsel 0 1 116			tpfbsel 0 1 116

tpfbsel 1 5 12c			tpfbsel 1 5 12c

testp 4e2 10f 1e			testp 4e2 10f 1e

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop He | ~600 |  |
| Stim evt cnt | ~600 |  |

### Stopping Z>2 Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 1 6				tpfbsel 0 1 6

tpfbsel 0 b 116			tpfbsel 0 a 116

tpfbsel 1 5 11c			tpfbsel 1 5 11c

testp 4e2 179 3d			testp 4e2 179 3d

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop heavies | ~600 |  |
| Stim evt cnt | ~600 |  |

### Penetrating H Events

Send the following HET commands:

FM1					FM2

tpfbsel 1 3 126			tpfbsel 1 2 126

tpfbsel 1 7 11c			tpfbsel 1 7 11c

tpfbsel 1 9 11c			tpfbsel 1 a 11c

tpfbsel 1 c 116			tpfbsel 1 d 116

testp 4e2 17 b				testp 4e2 16 a

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Pen que | ~600 |  |
| Pen H | ~600 |  |
| Stim evt cnt | ~600 |  |

### Stimulus Pulser / Analog Channels  (FM1 only, skip to step 6 for FM2)

First, send the following HET commands:

gwrite 9 32

tmode 3

hgorsel 1 0 12a8

lgorsel 1 0 12a8

The next section has been designed to pulse all the channels (H1i, H1o, H2, H3, H4, H5, H6).    The command set should be sent and the data values should be reviewed and recorded at each configuration as directed.  It should be noted that the expected pulse height value range may not be exact, but the values should lie very close to this range.

### Configure All Channels

Send the following HET command set to configure the channels as follows:

- enable the test pulser for all channels

- set the feedback capacitor to 30pF  for H1i, H1o, H2 and H6

- set the feedback capacitor to 60pF  for H3, H4 and H5

- set the test input capacitor to 18pF for H1i, H1o, H2 and H6

- set the test input capacitor to 36pF for H3, H4 and H5

tpfbsel 0 1 166

tpfbsel 0 b 166

tpfbsel 1 3 166

tpfbsel 1 5 1cc

tpfbsel 1 7 1cc

tpfbsel 1 9 1cc

tpfbsel 1 c 166

Send the following HET command set to configure the channels as follows:

- disable the high gain ADC for all channels

- enable the low gain ADC for all channels

phacont 0 1 7

phacont 0 b 7

phacont 1 3 7

phacont 1 5 7

phacont 1 7 7

phacont 1 9 7

phacont 1 c 7

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 5V

- set the Phasic 1 DAC setting to 5V

testp 4e2 1ff 1ff

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c813 – 08c818 |  |  |
| H1o | 05c80b – 05c80f |  |  |
| H2 | 11c804 – 11c808 |  |  |
| H3 | 12c830 – 12c834 |  |  |
| H4 | 13c818 – 13c81d |  |  |
| H5 | 14c824 – 14c828 |  |  |
| H6 | 96481c – 964821 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 2.51V

- set the Phasic 1 DAC setting to 2.51V

testp 4e2 180 180

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c423 – 08c425 |  |  |
| H1o | 05c41d – 05c420 |  |  |
| H2 | 11c41a – 11c41c |  |  |
| H3 | 12c42e – 12c430 |  |  |
| H4 | 13c422 – 13c425 |  |  |
| H5 | 14c428 – 14c42a |  |  |
| H6 | 964425 – 964427 |  |  |

Send the following command set to configure all the channels to the following:

- enable the high gain ADC for all channels

- enable the low gain ADC for all channels

- verbose mode off

phacont 0 1 d

phacont 0 b d

phacont 1 3 d

phacont 1 5 d

phacont 1 7 d

phacont 1 9 d

phacont 1 c d

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .25V

- set the Phasic 1 DAC setting to .25V

testp 4e2 ff ff

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0887e7 – 0887eb |  |  |
| H1o | 0587cb – 0587d0 |  |  |
| H2 | 1187eb – 1187f0 |  |  |
| H3 | 1287ec – 1287f0 |  |  |
| H4 | 1387e8 – 1387ed |  |  |
| H5 | 1487de – 1487e2 |  |  |
| H6 | 9607f6 – 9607fb |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 0.125V

- set the Phasic 1 DAC setting to 0.125V

testp  4e2 80 80

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08840b – 088410 |  |  |
| H1o | 0583fd – 058401 |  |  |
| H2 | 11840c – 118410 |  |  |
| H3 | 12840c – 128410 |  |  |
| H4 | 13840a – 13840f |  |  |
| H5 | 148405 – 148409 |  |  |
| H6 | 960411 – 960416 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .002V

- set the Phasic 1 DAC setting to .002V

testp 4e2 2 2

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08803d - 088040 |  |  |
| H1o | 05803b – 05803f |  |  |
| H2 | 11803b – 11803e |  |  |
| H3 | 12803a – 12803e |  |  |
| H4 | 13803b – 13803e |  |  |
| H5 | 14803a – 14803e |  |  |
| H6 | 96003a – 96003e |  |  |

Finally, send the following HET commands or reboot:

testp 4e2 0 0

gwrite 9 30

tmode 0

hgorsel 1 0 8

lgorsel 1 0 8

tpfbsel 0 1 6

tpfbsel 0 b 6

tpfbsel 1 3 6

tpfbsel 1 5 c

tpfbsel 1 7 c

tpfbsel 1 9 c

tpfbsel 1 c 6

testp auto

*** END OF TEST FOR FM1.  STOP HERE IF TESTING FM1.***

### Stimulus Pulser / Analog Channels  (FM2 only)

First, send the following HET commands:

gwrite 9 32

tmode 3

hgorsel 1 0 24a4

lgorsel 1 0 24a4

### Configure All Channels

Send the following HET command set to configure the channels as follows:

- enable the test pulser for all channels

- set the feedback capacitor to 30pF  for H1i, H1o, H2 and H6

- set the feedback capacitor to 60pF  for H3, H4 and H5

- set the test input capacitor to 18pF for H1i, H1o, H2 and H6

- set the test input capacitor to 36pF for H3, H4 and H5

tpfbsel 0 1 166

tpfbsel 0 a 166

tpfbsel 1 2 166

tpfbsel 1 5 1cc

tpfbsel 1 7 1cc

tpfbsel 1 a 1cc

tpfbsel 1 d 166

Send the following HET command set to configure the channels as follows:

- disable the high gain ADC for all channels

- enable the low gain ADC for all channels

phacont 0 1 7

phacont 0 a 7

phacont 1 2 7

phacont 1 5 7

phacont 1 7 7

phacont 1 a 7

phacont 1 d 7

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 5V

- set the Phasic 1 DAC setting to 4.71V

testp 4e2 1ff 1f0

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height)  using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c816 – 08c818 |  |  |
| H1o | 054808 – 05480d |  |  |
| H2 | 1147e8 – 1147ed |  |  |
| H3 | 12c7ea – 12c7ef |  |  |
| H4 | 13c7f0 – 13c7f5 |  |  |
| H5 | 1547ea – 1547ef |  |  |
| H6 | 96c7dc – 96c7e1 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 2.51V

- set the Phasic 1 DAC setting to 2.51V

testp 4e2 180 180

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c423 – 08c426 |  |  |
| H1o | 05441c– 05441e |  |  |
| H2 | 11444b – 11444e |  |  |
| H3 | 12c44c – 12c44e |  |  |
| H4 | 13c44f – 13c451 |  |  |
| H5 | 15444c – 15444e |  |  |
| H6 | 96c444 – 96c447 |  |  |

Send the following command set to configure all the channels to the following:

- enable the high gain ADC for all channels

- enable the low gain ADC for all channels

- verbose mode off

phacont 0 1 d

phacont 0 a d

phacont 1 2 d

phacont 1 5 d

phacont 1 7 d

phacont 1 a d

phacont 1 d d

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .25V

- set the Phasic 1 DAC setting to .235V

testp 4e2 ff f0

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0887cc – 0887d1 |  |  |
| H1o | 0507d2 – 0507d6 |  |  |
| H2 | 1107c9 – 1107cd |  |  |
| H3 | 1287ad – 1287b1 |  |  |
| H4 | 1387af – 1387b4 |  |  |
| H5 | 1507b4 – 1507b8 |  |  |
| H6 | 9687a3 – 9687a8 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .125V

- set the Phasic 1 DAC setting to .125V

testp  4e2 80 80

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0883fd – 088402 |  |  |
| H1o | 0503ff – 050404 |  |  |
| H2 | 110439 – 11043e |  |  |
| H3 | 12842a – 12842f |  |  |
| H4 | 13842b – 138430 |  |  |
| H5 | 15042e – 150432 |  |  |
| H6 | 968425 – 968429 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .002V

- set the Phasic 1 DAC setting to .002V

testp 4e2 2  2

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08803c – 08803f |  |  |
| H1o | 05003b – 05003e |  |  |
| H2 | 11003d – 110041 |  |  |
| H3 | 12803d – 128041 |  |  |
| H4 | 13803d – 138040 |  |  |
| H5 | 15003d – 150040 |  |  |
| H6 | 96803c – 96803f |  |  |

Finally, send the following HET commands or reboot:

testp 4e2 0 0

gwrite 9 30

tmode 0

hgorsel 1 0 4

lgorsel 1 0 4

tpfbsel 0 1 6

tpfbsel 0 a 6

tpfbsel 1 2 6

tpfbsel 1 5 c

tpfbsel 1 7 c

tpfbsel 1 a c

tpfbsel 1 d 6

testp auto

## 4.4 SIT  Comprehensive Test Procedure (In Flight)

Send the following SIT commands:

immed 1	causes all commands to be executed upon receipt

hvenable 0	ensures that the SIT HVPS is disabled

junk 1		causes SIT to respond to all events, even illegals

eonly 1	removes TOF coincidence requirement for events

toferror 1	processes events including TOF errors

limhi 100

Verify SIT is configured for quiet mode. After several minutes, verify the parameters from the display with the values in the following table.  These parameters will verify that SIT is in quiet mode.

:

| Description | Expected Value |
| --- | --- |
| Toferror | 1 |
| HV Status | 0 |
| Eonly | 1 |
| Junk | 1 |
| Limhi | 256 |

Verify SIT is configured in noisy mode

Send the following HAZARDOUS COMMANDS.

Double-check the commands before sending them.

hvenable 1	enables the SIT HVPS

hvlevel 10	sets the HVPS output to about 300V

Wait 2-3 minutes for the new telemetry and then verify the following status information has been returned on   SIT GSE and matches expected values:

| Description | Expected Value |
| --- | --- |
| HV Step | 10 |
| Toferror | 1 |
| HV status | 1 |
| Eonly | 1 |
| Junk | 1 |
| HV | 300-350 V |

After 5 minutes after beginning of SIT Noisy Mode send the following commands to turn off the HV.

hvenable 0	ensures that the SIT HVPS is disabled

hvlevel 0	sets HV level to 0

junk 0		causes instrument to reject illegal events

eonly 0	removes TOF coincidence requirement for events

toferror 0	processes events including TOF errors

limhi 1f4	limits high-priority events to 500/major frame

Wait 2-3 minutes for the new telemetry. In the meantime, continue with other Step 4 activities in parallel and verify the following status information has been returned on SIT GSE and matches expected values:

| Description | Expected Value |
| --- | --- |
| HV Step | 00 |
| Toferror | 0 |
| HV Status | 0 |
| Eonly | 0 |
| Junk | 0 |
| Limhi | 500 |

## 4.5 SIT High Voltage Power Supply Turn On Procedure (In Flight)

The SIT high voltage turn-on takes 5 days to reach the operating voltage of ~3000v.  This procedure assumes that the instrument will be powered on continuously for 5 consecutive days.

Verify SIT table checksums available in housekeeping data on SIT GSE:

| Description | Expected Value | Actual Value |
| --- | --- | --- |
| SIT FM1 tab cksm | 52A82E (4/28/06) |  |
| SIT FM2 tab cksm | 46E9B7 (4/28/06) |  |

Verify the following singles rates on SIT GSE match expected values (cts/min):

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | >600 |  |  |
| VSE | 0 |  |  |
| Matrix Rates (ApId 605) | All 0 |  |  |
| Beacon Rates (ApId 619) | All 0 |  |  |
| Pulse Height Events (ApId 606-617) | All 0 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected 				 values:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SW Ver | 0602 |  |  |
| SW error | 0 = (no error) |  |  |
| HV Step | 00 |  |  |
| Toferror | 0 = error events |  |  |
| HV Status | 0 = off |  |  |
| Eonly | 0 = ET coincidence required |  |  |
| Junk (Events) | 0 = junk events ignored |  |  |
| LIMHI | 500 |  |  |
| calib gain | 9 - 11 |  |  |
| calib offset | -14 to –65, should settle at –15 but may take as much as an hour (don’t wait for it!). |  |  |
| calib error | 0x08, should settle to 0x00 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected values:

| Description | Expected Values | Actual Values | Comments |
| --- | --- | --- | --- |
| HV | -79.6 |  |  |
| Tof T | 25-35 C |  |  |
| Foil T | 20-25 C |  |  |
| SSD T | 20-25 C |  |  |
| +3.3 V monitor | 3.3 V |  |  |
| +2.5 V monitor | 2.5 V |  |  |
| +5 V monitor | 5 V |  |  |
| +6 V monitor | 6 V |  |  |

Command sequence for HVPS turn on

**************************** Day 1 ****************************

Send the SIT Commands:

immed 1

hvenable 1

CAUTION: the following is a hazardous command.  Check carefully before sending

hvramp 40

After 2-3 minutes, verify the HV setting and count rates:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 40 |  |  |
| HV | 985 - 1045 |  |  |
| HV Status | 1 = on |  |  |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | >600 |  |  |
| VSE | 0 |  |  |

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp 80

After 2-3 minutes, verify the HV setting and count rates:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 80 |  |  |
| HV | 2063 - 2156 |  |  |
| HV Status | 1 = on |  |  |
| SRT | 0-10 |  |  |
| STP | 0-10 |  |  |
| VS | 0-10 |  |  |
| SSD | >600 |  |  |
| VSE | 0-10 |  |  |

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp 90

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 90 |  |  |
| HV | 2345 - 2422 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >10 |  |  |
| STP | >10 |  |  |
| VS | >0 |  |  |
| SSD | >600 |  |  |
| VSE | >0 |  |  |

**************************** Day 2 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp a0

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | A0 |  |  |
| HV | 2610 - 2704 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >10 |  |  |
| STP | >10 |  |  |
| VS | >0 |  |  |
| SSD | >600 |  |  |
| VSE | >0 |  |  |

**************************** Day 3 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp a8

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | A8 |  |  |
| HV | 2760 - 2853 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

**************************** Day 4****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp b0

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | B0 |  |  |
| HV | 2843 - 2986 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

**************************** Day 5 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully before sending

hvramp b8

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | B8 |  |  |
| HV | 2976 - 3109 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

******************************************************

Notes:

(1) range of expected HV values based on range between fm1 and fm2 hot/cold cases observed during thermal vac, ± 30 V allowance for 2 channel jitter in ADC

(2) range of expected count rates in detectors variable due to diferent levels of activity observed in IP space.  Figure below shows COUNTS/SECOND for Wind/STEP for comparable count rates during first portion of 2006