Building 510C

P.O. Box 5000

Upton, NY 11973-5000

Phone 516 344-6102

Fax 516 344-3253

sivertz@bnl.gov

managed by Brookhaven Science Associates
for the U.S. Department of Energy

date:		May 17, 2000

to:		Yousef Makdisi

from:		Michael Sivertz

subject:	Minutes of ORR and Walk Through of the PHENIX Gas Mixing House

The following represents the PHENIX response to the committee’s findings:

- All outstanding ESRC action items shall be closed. (Makdisi)

All outstanding action items have been closed.  The LCVS is on Emergency Power.

- Relevant operating procedures including operational limits are to be provided. (PHENIX)

Operating Procedures for the RICH, PC, DC, and TEC gas systems have been provided.  They include operational limits.

- The PHENIX installed safety systems should be certified, and interlocked. (Pendzick)

Done.  Blue Sheet Recertification has been accomplished.

- Signed drawings of the installed gas mixing systems should be provided. (PHENIX)

Done.  All Gas System Drawings have entered PHENIX Configuration Control.

- Label the gas vent lines and provide appropriate postings. (PHENIX)

Done.

- Bond the Vent lines with # 10 wire. (PHENIX)

Done.

- Label the gas lines in the trench, clean it, and install the proper covers. (PHENIX)

Done.

- The gas sensors should be included in a regular maintenance program. (PHENIX)

Done.

- Drawings of the electrical rewiring of the GMH should be provided. (PHENIX)

Gas Mixing House Schematic Drawing will be converted to AutoCAD.  These, and all other SMCS Drawings, will be formatted according to C-A Standards, given C-A Drawing Numbers and enter C-A Configuration Control.  This will be accomplished before May 31, 2000.  Point-to-Point Wiring Diagrams will be converted to C-A format, and enter C-A Configuration Control before July 15, 2000.  This represents a POST START Action Item.

- The DC/PC and TEC sampling compressors should be locked out. (PHENIX/ Pearson)

These compressors have been entirely removed from the system.  They will not be used as part of any PHENIX Gas Rack.

- Patch and seal the wall between the gas mixing room and the control room. (PHENIX)

Done.

- Clean and remove the construction debris. (PHENIX)

Done.

- The vent fan should be on a regular maintenance schedule. (PHENIX / Pearson)

Done.

- The ODH sensor should be operational and its sound alert reduced. (Williams)

Done.

- The electrical outlets should be changed to gfci. (PHENIX)

Done.

- Close up all holes in the gas mixing room and test the ventilation flow with smoke. (PHENIX / Pearson)

Done.

- Add a building number to GMH (PHENIX)

A temporary Building Number Sign has been affixed to the GMH.   A permanent sign is on order.

- Post a “keep closed” sign on the door between the control room and the gas mixing room. (PHENIX)

Done.

- Label all electrical panels and switches in the control room. (PHENIX)

Done.

- Label the normal trailer a/c power disconnect. (PHENIX)

Done.

- Update the fire department run books for the service disconnect locations. (Giannotti/ Levesque)

We are waiting for Plant Engineering to enter Building 1008F into their database.  When this has occurred, Paul Giannotti will update the run books for the fire department.  This represents a POST START Action Item.

- Anchor junction boxes inside the access panels in the control room. (PHENIX)

Done.

- Support the MC cable and replace the ceiling tiles. (PHENIX)

Done.

Please let me know if there are any issues I have not addressed.  I look forward to hearing from you as soon as we have been approved to flow flammable gas.

Sincerely,

Michael Sivertz

cc: 	Sam Aronson

Michael Marx

Bill Lenz

# Text boxes

Memo