GMP 11

Good Measurement Practice

for

Assignment and Adjustment of Calibration Intervals for Laboratory Standards

- Introduction

  - Purpose

Measurement processes are dynamic systems and often deteriorate with time or use.  The design of a calibration system would be incomplete without some established means of determining how often to calibrate instruments and standards.  A calibration performed only once establishes a one-time reference of uncertainty.  Recalibration detects uncertainty growth and serves to reset values while keeping a bound on the limits of errors.  A properly selected interval assures that an item will receive recalibration at the proper time. Proper calibration intervals allow specified confidence intervals to be selected and they support measurement traceability.  The following practice establishes calibration intervals for standards and instrumentation used in measurement processes.

  - Prerequisites

    - Calibration history for laboratory standards

    - Expected tolerance limits if applicable

  - Safety

    - No outstanding safety concerns

- Methodology

  - Summary

Recommended calibration intervals are based on various examples of traceability as described in GMP 13.  As data is collected and evaluated, the laboratory technical manager may adjust the calibration intervals to ensure that measurement results are not invalidated by the intervals selected.

  - Apparatus

None.

  - Procedure

    - Identification of Parameters

Critical Parameters

Components that contribute more than 25 % of a measurement’s uncertainty are identified as critical parameters.  To ensure an accurate evaluation of performance, calibration intervals are determined to meet a 99 % reliability target. Critical parameters are checked and defined below:

Mass Critical Parameters  (The example provided below corresponds to Option A in Mass Traceability in GMP 13.)

Balance Performance

Balance performance in Echelon 1 weighing processes is evaluated in every measurement series.  An F-test ratio evaluates the observed standard deviation of the process against the accepted standard deviation of the process.  The balance performance component is also tested with a check standard in each weighing series.  The check standard value is evaluated with a t-test by compared the observed value to the accepted reference value.  All other weighing processes have incorporated measurement control procedures and control charts that are evaluated as data is collected.

Mass Standards

Mass standards are dynamic with use.  Wear, contamination and other factors can cause drift from accepted values.  Thus, the following intervals have been set:

Table 1.  Calibration intervals for mass standards

| Item | Initial Cal Interval (months) | Source |
| --- | --- | --- |
| P1. kg + P1.. kg | 48 | NIST |
| C1. kg + C1.. kg | (alternating 2 years) 48 | NIST |
| P30 kg - P2 kg | 12 | Lab |
| P500 g - P1 mg | 6 | Lab |
| C500 g - C1 mg | 6 | Lab |
| W25 kg - W1 mg | 12 | Lab |
| P Pound Standards | 24 | Lab |
| W Pound Standards | 12 | Lab |
| P = primary; C = check/control; W = working standards |

### Length Critical Parameters

Length Comparison Performance

The measurement performance of each length calibration is evaluated with a check standard (when performing tape-to-tape comparison).

Initial Length Intervals

Length standards are dynamic with use.  Wear, contamination and other factors can cause drift from accepted values.  The following intervals have been set due to these factors:

Table 2.  Calibration intervals for length standards

| Item | Initial Cal Interval (months) | Source |
| --- | --- | --- |
| 100 ft Tape #1 | 60 | NIST |
| 100 ft Tape #2 | 60 | NIST |
| 25 ft or 7 m Tape | 60 | NIST |
| 18 in Steel Rule | 120 | NIST |
| Length Bench | 24 (if used or moved) | Lab |

Volume Critical Parameters  (Example shown corresponds to Option A for volume calibration in GMP 13.)

Volume Comparison Performance

The measurement performance of a volume transfer calibration is evaluated in each use with a repeatability check.  Use of check standards is the preferred method for evaluating the measurement process over time.  Traceability for volume standards may be established through gravimetric calibrations using traceable mass standards.

Initial Volume Intervals

Volume standards are dynamic with use.  Wear, contamination and other factors can cause drift from accepted values.  Calibration intervals are as follows:

## Table 3.  Calibration intervals for volume standards

| Item | Initial Cal Interval (months) | Source |
| --- | --- | --- |
| P100 gal standard | 60 | NIST |
| P25 gal standard * | 60 | Lab |
| P5 gal standard * | 12 | Lab |
| Glassware- Autopipetes 5 L to 100 mL | 120 | Lab |
| *Gravimetric calibration for volumes 5 gallon or smaller, and all “slicker plate” standards. Laboratory must be qualified for performing gravimetric calibrations. Volume transfer is acceptable above 5 gallon. *May be a “slicker plate” type. None are hand-held, “dump” style, test measures. |

Temperature Critical Parameters

Temperature Comparison Performance

The measurement performance of each temperature comparison calibration is evaluated with a check standard and can be verified periodically using triple point cells, melting point cells, and ice baths (using documented procedures).

Initial Intervals

Temperature standards are dynamic with use.  Shock, contamination and other factors can cause drift from accepted values.  Recalibration intervals are as follows:

Table 4.  Calibration intervals for temperature standards

| Item | Initial Cal Interval (months) | Source |
| --- | --- | --- |
| 25.5 ohm SPRT | 36 | NIST |
| 100 ohm PRT′s | 12 | Lab |
| Standard Thermistor | 12 | Lab |
| Check Standards | 6 | Lab |
| Liquid-in-glass standards | 6 | Lab |

Secondary Parameters

Components that contribute less than 25 % but more than 1 % of a measurement’s uncertainty are identified as secondary parameters. Secondary parameters are assigned calibration intervals designed to meet a 95 % reliability target.  Secondary parameters are defined below:

Mass and Gravimetric Volume Secondary Parameters

Environmental Measurement Equipment

Table 5.  Calibration intervals for environmental equipment

| Item | Initial Calibration Interval (months) | Source |
| --- | --- | --- |
| Barometer | 12 | NIST |
| Hygrometer | 24 | NIST |
| Thermometer | 12 | Lab |

Length Secondary Parameters - No secondary parameters.

Volume Secondary Parameters

Water Temperature Measuring Device

Table 6.  Calibration intervals for volume secondary parameters

| Item | Initial Cal Interval (months) | Source |
| --- | --- | --- |
| Thermometer | 12 | Lab |

Temperature Secondary Parameters - No secondary parameters.

- Calculations

  - Initial Intervals

    - Assignment of Initial Intervals

Assignment of initial intervals is based on these recommendations.  Subsequent intervals may be adjusted based on analysis of Check standard data that is recorded on control charts.  The initial interval is equivalent to the amount of time that passes before a metrologist makes the first observation of a result lying outside the warning limits of the chart when the investigation yields no apparent correctable cause.  Warning limits are established at ( 2 standard deviations of the measurement process around the accepted value of the check standard.

    - Absence of Control Charts

If no control charts are available, the laboratory’s Technical Manager will assign the initial interval based on engineering evidence, manufacturer(s specifications, NIST recommendations, and experience.

  - Adjustment of Intervals

    - Intervals will be adjusted when determined to be necessary by the laboratory’s Technical Manager.

    - The intervals will be adjusted by taking the following factors into consideration as appropriate:

      - calibration history;

      - measurement assurance data;

      - data for the population of similar equipment;

      - NIST recommendations;

      - statistical analysis methods; and

      - manufacturer(s recommendations.

- Assignment of Uncertainty

The uncertainty associated with the establishment of calibration intervals is not included as a part of the uncertainties associated with the respective measurement SOPs.  See SOP 29 for calculating uncertainties for standards and the instruments.