Tevatron Ion Profile Monitor (IPM) Project

IPM Front End (IFE)

Kwame Bowie

August 10, 2005

Revised: December 9, 2005

Introduction:

This document will present the current state of the multi-channel QIE board designed for the Tevatron IPM project. The motivation of this document is to describe the board in its present state. This document will discuss the board’s architecture, and its operation within the Tevatron IPM DAQ system.

Tev IPM DAQ System Architecture:

The IPM Front End (IFE) board is the front end portion of the IPM data acquisition (DAQ) system. The following shows in block diagram form the basic architecture and data flow of the DAQ system:

The IFE board is a single element of the Tev IPM DAQ system. The Tev IPM DAQ system may be thought of as two separate subsystems: upstairs and in-tunnel. The upstairs portion of the DAQ system all resides in a single PC. The upstairs PC houses the Timing Card as well as the Buffer Boards and implements all of the intelligence and control for the DAQ system. The in-tunnel portion of the DAQ system all resides in a single 6U x 160mmEurocard crate. The in-tunnel portion includes sixteen (16) front end cards as well as a single fanout card. The in-tunnel portion intentionally has very little memory or configurable functionality because it will sit very close to the beam line in a moderately high radiation dosage area. For that reason, all components of the IFE and fanout cards are designed to tolerate reasonable dosage levels, and unnecessary logic is not present in the design. The remaining sections of this document will discuss the IFE board design and functionality in detail.

IFE Specifications:

The following specifications have been developed as the result of several meetings with various members of the project group. The specifications remain somewhat malleable, until the full IPM system has been calibrated. This document breaks the specifications down into two categories: functional and performance.

Functional Description:

- Digitize the charge data of 8 QIEs operating in both calibration mode and normal non-inverting mode.

- Relay the following flags to closest QIE clock cycle:

  - Proton

  - Pbar

  - Proton Injection

  - Antiproton injection

  - Error bits

  - Crossing(clock) counter

  - Proton bunch counter

- Multiplex data for serialization via an optical link

- Integrate charge at rate of 2RF/7

Performance Requirements:

- Operation reliable up to 20 krad total ionizing dose (TID)

- Noise level: 0.5 – 1LSB RMS (1-2 fC)

- Minimum signal (resolution): ~1 fC

- Provide additional 4- 26 bits of header data for each integration cycle

- High speed serial data out using optical link @1.6 Gbps

IFE Board Architecture

The block diagram below illustrates the main functional elements of the IFE board. The following sections will provide more detail about each of the blocks.

QIE biasing and interface

The core element of the charge integration and digitzation for the IFE board is the QIE8 developed for CMS by Fermilab engineers. The principal designer T. Zimmerman has authored several documents available for review which provide a very detailed description of the operation and performance of the chip. The QIE is an amazing chip that can operate in many different configurations in order to support many different detector signal sources. In this section, we will discuss the configuration that we have selected for the QIE. The table below illustrates the all of the possible modes of operation of the QIE.

| Range mode | Charge Polarity | Input Impedance |
| --- | --- | --- |
| Calibration Mode | Non-inverting | 50 ohm |
| Normal Mode | Inverting | 90 ohm |

The QIE may operate in non-inverting mode or inverting mode. Inverting mode is designed to be used with positive current source detector elements such as HPDs. The IFE board is designed to run with the QIE operating in its non-inverting mode. The non-inverting mode of the QIE is useful for detector elements generating negative current pulses such as PMTs or similar devices. The micro channel plates (MCPs) collecting ionization electrons provide a similar type of pulse when configured for the IPM.

The QIE also has selectable input impedance. One may select between 90ohm or 50 ohm termination inside the QIE chip. The IFE board utilizes the 50ohm input termination for the QIEs amplifiers. The selection of the 50ohm termination was primarily due to cabling and connector concerns.

The QIE integrates charges with a dynamic range of roughly 14 bits in its NORMAL mode of operation. Although there are 14 bits of dynamic range, the data is condensed into a single 9-bit data word with non-linear transfer function. Although large dynamic range is quite useful in many applications, the signals of interest for the Tev IPM project reside in a fairly small window at the lowest end of measurable charges for the QIE. In fact, quite frequently the signals of interest are below the threshold of the LSB in NORMAL mode. The QIE has a CALIBRATION mode in which the voltage vs. charge curve is linear with a slope of  ~0.9fC/ count. In many situations, several channels per board will be expected to operate in NORMAL mode and the remainder in CALIBRATION mode. The IFE board provides a single 8-bit DIP switch between the two backplane connectors that enable switching to NORMAL mode. If the DIP switch is not set, then the QIE on the associated channel will always operate in CALIBRATION mode.

In addition to the mode settings, it is important to bias the QIE correctly for the signals of interest and the selected sampling rate. The QIE8 chip has a bias setting resistor that sets the bias currents for all of the internal transistors and amplifiers. By decreasing the bias, one may integrate over a longer period of time without saturating the output. A resistance of 750Kohms provides an ideal bias for the selected sampling rate of roughly 15MHz.

The QIE is capable of integrating and digitizing charge at a better than 40Mhz sampling rate. This high throughput may only be accomplished by using a multi-cycle pipeline. The QIE has a four-cycle pipeline with four identical and parallel charge integration pathways. Unfortunately, these identical pathways have slightly different characteristics due to wafer variations that simply cannot be perfectly matched and controlled. Since these pathway-to-pathway variations will remain nearly constant over time, a great increase in accuracy may be extracted by calibrating each of these pathways. In order to calibrate these pathways, we must know which pathway each sample represents. The QIE provides this pathway information in the form of two additional bits for each sample called CAPID bits. Without these CAPID bits and a full calibration, the QIE charge measurement accuracy is of order 25% while with a full calibration, the accuracy of 5% or better is achievable.

Timing/Control Link:

The timing/control link originates from the timing card in the upstairs subsystem. The timing/control link signals are sent from a PECL driver in the upstairs subsystem to the in-tunnel subsystem on a single CAT5 cable. Once the signals reach the tunnel, the signals are duplicated and passed to the IFE boards by the Fanout Card on custom modular cables. The timing/control link of each IFE is composed of 4 beam-synchronized PECL pairs signals. The signals are listed below:

  - CLOCK (2RF/7)

  - PROTON marker

  - PBAR (antiproton) marker

  - CONTROL (5-bit serial encoded)

The PROTON and PBAR markers are provided by the timing card to ease the load on the Buffer Board triggering logic. The Tevatron operates with a known beam structure and bunch spacing. Relevant charge signals are only expected during the known bunch crossings. Providing beam-synchronized PROTON and PBAR signals to the Buffer Board allows these signals to used as a trigger for storing data. Without the PROTON and PBAR signals, the Buffer Board would need to trigger on data in a possibly very complex manner. The two markers provide a very accurate and very simple trigger for data storage.

Clock Distribution:

The clock distribution on the IFE board includes two (2) asynchronous clock domains: the system or QIE clock, and the serializer clock. The distribution of each of these clocks to the relevant portions of the board is discussed in this section

The clock distribution scheme for the QIE clock is based on a tree of point-to-point links as shown in the diagram. The fanout card distributes beam timing information to each front end via a dedicated point-to-point link. Each IFE board then duplicates the system clock signal nine (9) additional times and sends the clock signal to the eight (8) QIEs  and an FPGA via point-to-point links. The total skew between all clocks is less than +/- 2 nanoseconds.

The GOL serializer requires a very pristine clock signal. The GOL expects a clock with less than 150 ps of jitter peak-to-peak. When designing the clock distribution, each layer of logic adds a small amount of jitter to a clock signal. The GOL’s strict clock jitter requirements suggested that minimum logic be placed between the origin of the serializer clock and the GOL chip. The simplest solution turned out to be placing a low-jitter clock on each board with a point to point link to the GOL.

Data Serialization

The IFE board uses the GOL as its serializing circuit. The GOL is an ASIC developed for the CMS project at CERN. The GOL has multiple operating modes, but we will operate constantly in the high speed mode allowing serialization at a bit rate of 1.6 Gbps. The GOL may be thought of as a high speed parallel-to-serial shift register that receives a 32-bit input and transmits a 1.6Gbps output.

As the IFE board and DAQ system will be based upon two asynchronous clocks, it is very important to understand the relationships between these two clocks and its impact on the data. The two clocks that must be present for the system to operate reliably are the system clock and the serializer clock.

The system clock is based upon the tevatron RF cycle. The frequency of the system clock is actually two-sevenths of the tevatron RF frequency or 15.17 MHz. This beam-synchronized clock provides us our time reference point for our signals of interest. Unfortunately the frequency of this clock varies slightly with the Tevatron’s mode of operation. This variation in the system clock makes it unsuitable for use as the serializer clock due to its frequency variation.

To enable reliable serialization, a separate high-precision crystal is located on the IFE. The frequency of this crystal is defined by the bit-rate of the serializer and for our purposes is a factor 40 times smaller than the bit rate of serialization. The GOL ASIC accepts a 40MHz clock from these crystals and serializes data at the bit rate of 1.6Gbps.

Since this system is based on two asynchronous clocks that are not pure harmonics of one another, there is a definite need to ensure that data is not lost or re-transmitted. Either situation makes the data acquisition a more complex problem. To make the receiving of the serial data the simplest, only relevant data will be transmitted. Any stale data will not be transmitted during its respective serializer clock cycle. Instead an idle (K28.5) character will be sent, which will notify the receiver that the previous word was not valid data. One requirement that will be made is that there must be a fixed-length data frame format. This simplifies the encoding and decoding processes by not requiring large amounts of memory and decision circuitry. The fixed length data frames may be separated by IDLE characters of arbitrary length.

Over time the two clocks will be related in a semi-periodic way, and this relationship will be discussed below:

Digitizing 8 QIEs @ 1.6 Gbps using GOL ASIC:

For an 8-channel board using the CERN GOL chip for serialization at 1.6Gbps, the serializer clock would be 40 MHz. The relationship between these two clocks is such that it will take several cycles of the system (QIE) clock for there to be an extra serializer clock cycle. The calculations based upon ideal clocks (no frequency variation) would look as follows:

Clock period relationship:

System clock period =  1 / 15 MHz = 65.908 ns

Serializer clock period = 1 / 40 MHz = 25 ns

Ratio of Serializer clocks per system clock period =  2.64:1

Data bus width relationship

QIE data bits generated each QIE clock = 58 bits (keeping only 1 CAPID)

GOL data bus width = 32 bits

Ratio of bits transferred per clock cycle =  1.81

These two relationships must be used when designing a fixed-length data frame format. The clock period ratio of 2.64:1 defines the fixed clock relationship and as a result, the maximum ratio of output (serializer clock) data words-to-input (qie clock) data words. This ratio of course refers to integer number of input data words and output data words. It is possible to transfer less information than this ratio by utilizing IDLE characters. However, it is beneficial to utilize as much bandwidth as possible. The data per cycle ratio of 1.81 defines the minimum ratio of serializer words to QIE words to ensure that there is not an overflow and loss of relevant data. There are really three tradeoffs that must be balanced when designing the data format:

- Bandwidth utilization

- Ease of encoding/decoding

- Information for each QIE clock cycle

The two ratios define the upper and lower bounds of the ratio of serializer clock cycles per qie clock cycle of the fixed length data frames. The following examples will demonstrate some of the tradeoffs.

Single QIE clock cycle design:

The use of a single QIE clock cycle means that only integer ratio values are achieveable. As a result, a ratio of 2 provides the only solution that lies within the ratio boundaries. Thus the data frame has a length of 2 and it transmitted after every QIE clock cycle. The amount of data sent in this case is 64 bits, so 6 additional bits are provided every data frame. However, the ratio of 2 means that a significant portion of the time, IDLE characters are being sent and bandwidth is being wasted. Effectively, on average 2*2RF/7*32bits/cycle = 971 Mbps are being transmitted as the data rate. This data rate amounts to only 76% bandwidth utilization.

2 QIE clock cycle design

The use of 2 QIE clock cycles provides two acceptable solutions: a 4 word data frame and a 5 word data frame. These two solutions provide ratios of 2 and 2.5 respectively. The second solution is much preferred as it provides the maximum in terms of bandwidth utilization and extra data per QIE clock cycle. Using a 5 word frame, the amount of extra bits per QIE clock is as follows:

Extra bits  per QIE clock  =

Extra bits  = 22 per QIE clock cycle

Using the 5-word frame also yields a high bandwidth utilization of 1.21 Gbps of a possible 1.280 Gbps or 94 % utilization.

As you can see, using a 5 word frame based upon 2 QIE clock cycles provides a good solution. An even better utilization of bandwidth may be achieved by utilizing a 13 word frame based upon 5 QIE clock cycles (2.6 clock ratio, 98% bandwidth utilization, 25.2 bits per QIE clock). However, as the data frame grows larger, more memory is required to build and decode the frame and a more complex state machine is required to control the frame builder and the frame decoder. In addition, the decoder on the buffer board requires more memory and logic.

The frame length that the IFE currently implements is a 7-word frame based upon 3 QIE clock cycles. This solution provides for frequent IDLE characters, a simple encoder and decoder, and low memory requirements. The data frame format is shown in the diagram below.

Data Synchronization:

The IFE board accepts charge signals for eight (8) channels as its input and generates a high speed data stream as its output. These signals are always running and are presented to the receiver boards without interruption. The IFE board from a functionality perspective is simply a time multiplexing and digitizing board. There is very little memory and no trigger generation currently present in the IFE board design. It is critical that IFE board data bytes may be associated with the correct Tevatron events. To ensure that data can be correlated with Tev events, synchronization must be accomplished.

There are essentially two requirements to achieve synchronization between the different IFE boards in the Tev IPM DAQ system. The first requirement is at least one common time reference. The timing control link provides the common clock for all IFE boards. The second synchronization requirement is that there must be a synchronization input. The synchronization input is implemented with the CONTROL signal within the timing/control link. The RESET command serves as the synchronization command as it ensures that all QIEs begin integration and transmission during the same Tevatron event.

Although all IFE boards begin integration and transmission during the same Tevatron event, the data links on each IFE board will operate at slightly different data rates. The key element to the synchronization is the FIFO internal to each IFE and ability of the serializer to send IDLE characters.

Since the actual data throughput is much less than the link data rate, it is possible to transmit IDLE characters where stale data would be. The Buffer Board then stores only the valid data and not the IDLE characters to memory. The differing data link rates for the different IFE boards means that some boards will have more IDLE characters, but all boards will have the same amount of valid data. If only the valid data is stored to memory and the first word from each IFE board is aligned, and no data is lost, then  the data is guaranteed to be aligned and synchronization is achieved. The main challenge of this synchronization technique is maintaining a solid lock on the data link thus never losing data.

The Buffer Board must align the data from all IFE boards to synchronize the boards. To align the data from the boards, there must be a known time reference. There are several markers as well as several counters that help accomplish this task.

.

FPGA Design:

The FPGA design for the IFE board has two key functions: mode control and multiplexing/frame building. The FPGA has a few other responsibilities, but these two are the key components of the FPGA design.

The mode control portion of the FPGA design simply decodes the command sent on the CONTROL signal from the timing/control link. The control signal has available 8 different commands. The commands are encoded such that only commands starting with a “11” will trigger the decoding circuitry. The table below lists the commands that the IFE board implements and the associated codes that must be present on the CONTROL signal in five (5) consecutive system clock cycles.

Code			Function

| 11000 | No effect |
| --- | --- |
| 11001 | Proton Injection Marker |
| 11010 | Antiproton Injection Marker |
| 11011 | QIE Mode0 (QIE in Calibration Mode) |
| 11100 | QIE Mode1 (QIE in Normal Mode) |
| 11101 | QIE Mode2 (Link test - outputs a 32-bit counter to GOL) |
| 11110 | QIE Mode3 (Reset QIEs) |
| 11111 | IFE Board Reset |

The functions are implemented such that there must be a one-cycle waiting period between sending back-to-back commands.

The other function of the FPGA is to implement the multiplexing and frame building components scheme. The implementation of the multiplexer and frame builder is accomplished by creating a high speed multiplexer that feeds a FIFO. The high speed multiplexer generates a 7-word frame after every 3 QIE clock cycles. This 7-word frame is then written into a FIFO. The FIFO directly feeds the GOL. If the FIFO empties, then an IDLE character is forwarded by the GOL. The interspersing of IDLE characters ensures that stale data is never sent and the Buffer Board only keeps valid data from each IFE board.

Power Distribution:

|  | Power per chip | Chips on board | Total power |
| --- | --- | --- | --- |
| CMS QIE8 | 600 mW @5 V | 8 | 4.8 Watts @ 5V |
| Serializer (TLK2501/GOL) | 360 mW @2.5V | 1 | 0.360 Watts @2.5V |
| VCSEL | 37.5mW @2.5V | 1 | 0.0375 Watts @ 2.5V |
| FPGA | 700 mW @3.3V, 200mW @ 1.5 | 1 | 0.900 Watts @3.3V, 1.5V |
| PECL Receivers | 110 mW @3.3V | 36 | 3.96 Watts @3.3V |
| LVDS-PECL terminations | 36mW @ 3.3V | 36 | 1.3068Watts @ 3.3V |
| PECL-PECL terminations | 52 mW @ 3.3V | 12 | 0.622 Watts@ 3.3V |
| PECL Clock fanout (MC100LVEP111) | 400 mW @3.3V | 1 | 0.400 Watts @ 3.3V |

12.4Watts per board

The power consumption estimates in the table above are listed in wattage. The values are listed in wattage to account for the varying operating voltages of all of the chips. Listing the values in wattages also enables the selection of an acceptable power supply. Different core voltages will be converted using linear regulators to step down the voltage from a higher voltage power supply. It is desirable to send a moderately higher voltage from the power supply to the IFE board. However, the linear regulators used by the IFE boards must dissipate internally any excess voltage provided to them. A good, safe assumption for the dropout voltage of a linear regulator is 1.5V. This means that the power supply that provides the input for the QIE regulator should be at least 7.0V and the regulator that provides the power for the remainder of the board should be at least 4.8V.

The IFE board sits in a 6U x160 Eurocard chassis with a custom backplane that provides two power buses per four IFE boards. It is expected that a custom rad-tolerant supply will be designed to provide power for the 16 IFE boards within a front end crate. The ideal supply would be able to generate the two voltages below at the corresponding currents:

QIE Supply:  7Volts at 20 Amps

3.3V-generating supply: 4.8 Volts at 40 Amps

The values above were rounded up to the nearest 10 amps to allow for dose-related current increases.