## Appendix A
ACTIVITY HAZARD DOCUMENT (AHD)  #2086

| GENERAL INFORMATION |
| --- |
| TITLE | LOCATION Bldg.-Room | DIVISION | LASER SAFETY CONTACT |
| Activity Hazard Document for the Laser Heated Diamond Anvil Cell Experiment at Endstation 2 of ALS Beamline 12.2.2 | Building 6 (Beamline 12.2.2) | Advanced Light Source | Sander Caldwell |

DESCRIPTION OF ACTIVITY

This AHD covers the Laser Heating Diamond Anvil Cell (LHDAC) at the ALS Beamline 12.2.2. The laser system is based on a commercial Quantronix Falcon –217 D Nd:YLF infrared (1053 nm) laser with a peak power of 80W in multimode in continuous wave (cw) mode. The laser head is on an optical bench approximately 3 feet above the hutch floor. The laser beam is directed to another optical bench above the first approximately 6 feet above the floor where it is split and directed downwards to the sample (LHDAC) which sits in the x-ray beam path. The laser system will be used to heat samples that are contained in a high-pressure apparatus so that the sample is simultaneously at high pressure and high temperature. The x-ray beam of beamline 12.2.2 will then be used to probe the crystal structure of the sample.

DURATION

X Ongoing

IDENTIFICATION OF HAZARDS Identification of potential hazards associated with the activity, lasers, and any additional hazards (e.g., toxic gases)

The hazard associated with the activity covered by this AHD is exposure to Class IV laser beams.

MITIGATION OF HAZARDS

Identification of laser(s): Laser specifications

|  | Laser 1 | Laser 2 | Laser 3 |
| --- | --- | --- | --- |
| Type | Nd:YLF infrared laser |  |  |
| Manufacturer | Quantronix |  |  |
| Model | Falcon 217D |  |  |
| Serial # | 0104-00314 |  |  |
| Maximum Power | 80W |  |  |
| Wavelength(s) | 1053 nm |  |  |
| Wavelength Used | 1053 nm |  |  |
| Power Used | 80 W |  |  |
| Pulse Length | continuous |  |  |
| Pulse Repetition Rate | continuous |  |  |
| Beam Diameter | 1 mm |  |  |
| Beam Divergence |  |  |  |
| Property # | 6481129 |  |  |
| Class | IV |  |  |

| LASER USERS |
| --- |
| Employee | Date Laser Safety Training Completed | Date Eye Examination Completed | Signature |
| Sander Caldwell | July 7, 2004 | July 8, 2004 |  |
| Martin Kunz | August 17, 2004 | Nov 15,2003 |  |
| Simon Clark | 01/06/04 | August 19,2004 |  |

ACCESS CONTROLS, INCLUDING USE OF INTERLOCKS

Access to the laser area is restricted as it is located at the Advanced Light Source. The laser area is entirely located within the hutch at beamline 12.2.2. The entrance to the hutch is marked with a laser warning sign identifying the area as containing a Class IV laser. An external shutter is mounted directly after the opening to the laser cavity and interlocked to the hutch door. Beam enclosures primarily consist of the existing x-ray hutch, whose one window has been covered with a special acryclic with an OD of >5 at 1053 nm. This essentially makes the laser system Class I during normal operation. Inside the hutch there is opaque aluminum paneling surrounding as much of the beam path as possible. The permanent beam enclosure will ensure that the entire top optic bench be enclosed by opaque aluminum panels, as well as the laser beam path as it makes its way from the lower bench to the upper bench. The sample area will also be enclosed by the opaque aluminum material. The top and side panels will only be removed during installation and alignment and will only be removable by deliberate use of tools (allen keys). Transport of laser beams between the upper and lower table is via opaque beam tubes. Beam tubes and enclosures will be marked with laser warning signs. Removal of enclosures, beam tubes, or panels will require tools to remove fasteners or keys to remove locks. All hazardous reflections will be terminated in beam stops at the end of their useful paths. Class IV laser warning signs will be posted on the laser and enclosures. Appropriate safety glasses will be worn during alignment procedures, and whenever Class IV laser beams are exposed.

ALIGNMENT PROCEDURE Procedural Considerations To reduce accidental reflections, remove watches, rings, dangling badges, necklaces, or reflective jewelry before any alignment activities begin. Use of nonreflective tools should be considered. Access to the hutch is limited to authorized personnel only. Consider having someone present to help with the alignment. Ensure that all equipment and materials needed are present prior to beginning the alignment. Ensure that all unnecessary equipment, tools, and combustible material (if fire is a possibility) have been removed to minimize the possibility of stray reflections and non-beam accidents. Persons conducting the alignment have been authorized by the responsible individual. Post a "Notice" sign at entrances when temporary laser control areas are set up or unusual conditions warrant that additional hazard information be available to personnel wishing to enter the area. Alignment Methods There shall be no intentional intrabeam viewing with the eye. Reduce the beam power by reducing power at the power supply. Avoid the use of high-power settings during alignment as much as is practical. Wear laser protective eyewear at all times during the alignment. Beam Control: The beam is enclosed as much as practical, the shutter is closed as much as practical during course adjustments, optics/optics mounts are secured to the table as much as practical, and beam stops are secured to the table or optics mounts. Areas where the beam leaves the horizontal plane shall be labeled. Any stray or unused beams are terminated. Invisible beams are viewed with IR cards, or similar technique. Operators are aware that specular reflections off some of these devices are possible, and that they may smoke or burn. No intrabeam viewing is allowed. Cameras or fluorescent devices are to be used for intrabeam viewing instead Normal laser hazard controls shall be restored when the alignment is completed. This includes enclosures, covers, replacement of beam blocks/barriers, and affected interlocks checked for proper operation. Beamline 12.2.2 Specific Alignment Procedures Only authorized laser users listed in this AHD may perform alignment and maintenance on the laser and optical systems covered by this AHD. The door to the hutch must be fully closed. All panels and optical shields must be in place and secured/locked properly during normal operation. Removing or opening the laser containment is only allowed for limited duration during beam alignment and maintenance. A sign will be posted on the hutch door indicating “DANGER-Visible and/or Invisible Laser Radiation-Avoid Eye or Skin Exposure to Direct or Scattered Radiation-Eye Protection Required”

Laser warning signs shall be posted on the hutch door and on the exterior of all beam enclosures.

EYE PROTECTION

It is Berkeley Lab policy that laser protective eyewear must be worn by all persons within the Nominal Hazard Zone (NHZ) whenever there is an open Class 3B or Class 4 beam. The NHZ is considered to be the entire room or laboratory in which the laser is located. The eye protection must have the appropriate optical density and/or reflective properties based on the wavelengths of the beams encountered, the beam intensity, and the expected exposure conditions. At the same time, the need for laser eye protection must be balanced by the need for adequate visible light transmission. Contact the LSO to select laser eye protection. Laser eye protection should be inspected periodically to ensure that it is in good condition.

| LASER PROTECTIVE EYEWEAR |
| --- |
| Number of Pairs Available | Location of Eyewear | Manufacturer | Optical Density | Wavelength |
| 2 | Inside Hutch | GPT | >7 | 850-1080nm |
| 1 | Inside Hutch | Laser Med | >4.5 | 1064 nm |
| 1 | Inside Hutch | Laser Peripherals | 5 | 1064 nm |
| 1 | Inside Hutch | Laser Vision | >5 | 1064 nm |
| 3 | Inside Hutch | GPT | 10 | 1064 |

BEAM PATH

The beam path is enclosed, as described above, except during commissioning and alignment. Procedures for commissioning and alignment are given above as well.

MAINTENANCE

The equipment will be maintained by specially trained and/or certified laboratory personnel. All relief devices, safety interlocks, alarms, and other hazard prevention devices will be maintained, calibrated, and tested for functionality on a regular basis in accordance with standard industrial practices and recommendations of the manufacturers.

EMERGENCY PROCEDURES

Authorized laser users will be familiar with the Building Emergency Plan, location of emergency equipment, and emergency procedures for fires, earthquakes, and evacuations. Emergency shut-off procedures for lasers consist of shutting off the electrical power to the laser system..

ANNUAL REVIEW SCHEDULE

Will be one year from approval date. If new hazards have been introduced, a full EH&S review will be required. If no changes other than users have been made (an update of the users list will be sent to LSO), renewal can be granted by user’s division safety coordinator.

## Berkeley Lab Incident Reporting Policy (Class 3B and Class 4 Laser Users)

Accidental laser exposure is a serious event that must result in ceasing operations pending investigation. The laser setup must be left unchanged to allow for the analysis of the cause of the accident. Exposed employees should report to Health Services in Building 26 to be evaluated for hidden injuries.