# Solicitation 1-057-RCA.1372  MICRO TESTER

# Performance Specification for MicroTester

The instrument must perform tensile, compression, flex, shear, cyclic and fatigue tests (both static and dynamic) on a variety of materials including fine wires and fibers (a few micron in diameter).

- The instrument must perform die shear test for semiconductor and button shear testing of molding compound.

- The instrument must perform monotonic and clyclic indentation tests.

- The instrument must perform peel test of thin films ans substrates and lap shear test of underfills and fracture toughness test of adhesions and interfaces.

- The instrument should be a compact size (500mm deep by 500mm wide) to be installed on any lab workbench or table.

## Frame

- It should be possible for the operator to configure the frame to be either vertical or horizontal (e.g. vertical for tensile testing on flexible specimens, horizontal for die shear).

- The instrument must employ a high-resolution linear optical grating transducer located directly on the actuator. The displacement accuracy must be better than  one micron and a mesurement resolution must be 20nm or better.

- The instrument must have a capability to employ a 500g load cell to measure down to 1.0gf with the accuracy within 1%.

- Maximum actuator travel and vertical test space must at least 100mm and 600mm, respectively.

- The drive system of the frame must employ a High Torque low inertia Brushless DC motor for excellent low speed smoothness and dynamic performance. Testing speed range must be at least 0.001 to 1,500 mm/min. The control loop should be isolated from the specimen via a ball screw reduction drive.

- The instrument must operate from a normal single-phase electric supply. No compressed air or hydraulic supply should be required.

- A fixed position emergency stop should be provided on the frame.

- Push buttons for coarse actuator positioning and a thumb wheel for fine actuator positioning must be provided at the frame.

## Controller

- The instrument must provide a high speed computer interface with a data transfer rate between the controller and PC of at least 7.7 Mbytes/sec continuously.

- The instrument must have a data sampling rate at least 5kHz on all transducer channels; data acquisition on all channels must synchronise to avoid data skew.

- The data resolution must be at least 19 bits per channel (1 part in 524,288) across the full scale of all transducers (i.e. no ranging).

- The Waveform generation should be at least 32bit resolution (1 part in 109).

- The controller must have at least one analogue input for an analogue waveform drive signal and at least four assignable analogue outputs.

- Transducer automatic recognition and calibration. The controller must have the facility to recognize virtually any number of transducers once they are fitted with the correct recognition/calibration plug without having to reconfigure either the controller or transducer.

- The controller must prevent any test proceeding with an uncalibrated transducer to prevent incorrect data being gathered.

## Software

- If a PC is used as an interface to the system the on-screen control display provided must always be clearly visible; it should not be possible to minimise or obscure this area of the screen.

- The control system must support, and the relevant software packages should be available to enable, the user to write and run their own application packages directly within a National Instruments Labview environment, without any additional compiling being required.

- Any software supplied must be able to run under Windows 95,98, NT and 2000.